# Ultraview Spec Sheet

Contents

1	Microscope	1

2	Objectives	1

2.1	Objective list	1

2.2	Lateral resolution	1

2.3	Axial resolution	1

3	Lasers	1

4	Filters	1

5	Cameras	1

6	Software	1

# Microscope

- PerkinElmer UltraVIEW Vox FRAP

- Body: Nikon Ti-Eclipse

- Spinning disk unit: Yokogawa CSU-X1

# Objectives

## Objective list

| Magnification | 100x | 60x | 40x | 20x |
| --- | --- | --- | --- | --- |
| Medium | Oil | Oil | Oil | Air |
| Brand | Nikon | Nikon | Nikon | Nikon |
| Type | Plan Apo VC | Plan Apo VC | Plan Fluor | Plan Apo VC |
| NA | 1.40 | 1.40 | 1.30 | 0.75 |
| WD | 0.13 | 0.13 | 0.20 | 1.00 |
| Coverslip | 0.17 | 0.17 | 0.17 | 0.17 |
| Correction collar | n | n | n | n |
| Spring | yes | yes | yes | yes |
| Tube | ∞ | ∞ | ∞ | ∞ |
| Bright field | DIC N2 | DIC N2 | DIC H | DIC N2 |

## Lateral resolution

- λem = 515 nm

- n = 1.518

- chip pixel size = 6.45 µm

| Objective mag. | NA | Pixel size (nm) | Rayleigh limit (nm) | Nyquist pixel size (nm) | Sampled? | Sampled w/ 1.5? |
| --- | --- | --- | --- | --- | --- | --- |
| 20 | 0.75 | 322.5 | 418.9 | 182.1 | No | No |
| 40 | 1.3 | 161.3 | 241.7 | 105.1 | No | No |
| 60 | 1.4 | 107.5 | 224.4 | 97.6 | No | Yes |
| 100 | 1.4 | 64.5 | 224.4 | 97.6 | Yes | Yes |

## Axial resolution

- λem = 515 nm

- n = 1.518

- physical pinhole radius = 25 µm

| Objective mag. | NA | Back projected pinhole size (um) | AU | z FWHM (nm, thin sample approximation) | z FWHM (nm, thick sample approximation) |
| --- | --- | --- | --- | --- | --- |
| 20 | 0.75 | 1.25 | 2.98 | 5478.7 | 7512.3 |
| 40 | 1.3 | 0.63 | 2.59 | 1303.2 | 2154.5 |
| 60 | 1.4 | 0.42 | 1.86 | 781.5 | 1367.4 |
| 100 | 1.4 | 0.25 | 1.11 | 554.7 | 908.1 |

# Lasers

| Wavelength (nm) | Output power (mW) | Imaging? | FRAP? |
| --- | --- | --- | --- |
| 405 | 50 | Yes | Yes |
| 488 | 50 | Yes | Yes |
| 561 | 50 | Yes | Yes |
| 640 | 50 | Yes | Yes |

# Filters

| Type | Device | Name | Ref No |
| --- | --- | --- | --- |
| Dichroic | CSU primary dichr | 405, 488, 561, 640 | 1 |
| Dichroic | CSU primary dichr | 405/488/561 | 2 |
| Dichroic | CSU primary dichr | 405-440/515/561 | 3 |
| Other | Em wheel 1 | Empty | 1 |
| Emission | Em wheel 1 | 525/50 | 2 |
| Emission | Em wheel 1 | 445/60, 615/70 | 3 |
| Emission | Em wheel 1 | 485/60, 705/90 | 4 |
| Emission | Em wheel 1 | 550/49 | 5 |
| Other | Em wheel 1 | Empty | 6 |
| Emission | Em wheel 1 | 446/32, 524/42, 600/36, 677/28 | 7 |
| Emission | Em wheel 1 | 478/45, 575/97, 707/86 | 8 |
| Emission | Em wheel 1 | 480/49, 536/24, 660/180 | 9 |
| Emission | Em wheel 1 | 448/55, 523/47 | 10 |
| Other | Em wheel 2 | Empty | 1 |
| Emission | Em wheel 2 | 590 LP | 2 |
| Emission | Em wheel 2 | Empty | 3 |
| Emission | Em wheel 2 | 485/60, 705/90 | 4 |
| Emission | Em wheel 2 | 575 LP | 5 |
| Emission | Em wheel 2 | 445/60, 615/70 | 6 |
| Emission | Em wheel 2 | 446/32, 524/42, 600/36, 677/28 | 7 |
| Emission | Em wheel 2 | Empty | 8 |
| Emission | Em wheel 2 | 480/49, 536/24, 660/180 | 9 |
| Emission | Em wheel 2 | 594/42, 730/140 | 10 |
| Dichroic | Storage | CSU prim dichr 405-445, 514-516, 638-647 | n/a |
| Emission | Storage | 525/50 | n/a |
| Emission | Storage | 550/49 | n/a |
| Other | Storage | DIC analyser (x2) | n/a |
| Emission | Storage | 478/45, 575/97, 707/86 | n/a |
| Emission | Storage | 457/50 | n/a |
| Dichroic | Ti base dichr wheel | Dapi Ex380 Em 440 | 2 |
| Dichroic | Ti base dichr wheel | GFP Ex485 Em515 | 3 |
| Dichroic | Ti base dichr wheel | mCherry (TRITC) Ex590 Em650 | 4 |
| Other | Ti base dichr wheel | empty | 1 |
| Other | Ti base dichr wheel | empty | 5 |
| Other | Ti base dichr wheel | empty | 6 |
| Dichroic | UltraVIEW sec dichr | CSU-X1 580 nm edge |  |

# Cameras

- 2 x Hamamatsu Orca-R2 interline CCD

- Speed = 16.2 fps (full-chip), 115.1 fps (8x8 sub-array, 8x8 binning)

- QE = 70 %

- Read noise (r.m.s) = 6 (slow mode), 10 (fast mode) electrons

- Dark current = 0.0005 electrons/px/s (at -40 °C)

## Field of view

- chip size = 1024 x 1344 px

- spinning disk aperture = 7 x 10 mm

| Objective mag. | Magnified chip size (µm) |  | Chip illumination area (µm) |  | FOV (µm) |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | x | y | x | y | x | y |
| 20 | 330.2 | 433.4 | 350.0 | 500.0 | 330.2 | 433.4 |
| 40 | 165.1 | 216.7 | 175.0 | 250.0 | 165.1 | 216.7 |
| 60 | 110.1 | 144.5 | 116.7 | 166.7 | 110.1 | 144.5 |
| 100 | 66.0 | 86.7 | 70.0 | 100.0 | 66.0 | 86.7 |

# Computer

- Software: Volocity v. 6.3.1

- OS: Windows 7 Professional 94-bit (6.1, build 7601)

- Sys manufacturer: Dell Inc

- Sys model: Precision Workstation T5500 (serial: BHL1X4J)

- BIOS: Phoenix ROM BIOS PLUS Version 1.10 A06

- Processor: Intel(R) Xeon(R) CPU E5640 @ 2.67GHz (4 CPUs), ~2.7GHz

- Memory: 4096MB RAM (DDR3)

Calculated using equation 26 in Wilson, T. Resolution and optical sectioning in the confocal microscope. Journal of Microscopy. 244:2 pp. 113-121 (2011)