# DV1 Spec Sheet

Contents

1	Microscope	1

2	Objectives	1

3	LEDs	2

4	Lasers	3

5	Filters	3

6	Cameras	4

7	Computer	4

# Microscope

- API Deltavision Elite DV1

- Body: Olympus IX-71

# Objectives

## Objective list

| Magnification | 100x | 60x (on request) | 40x | 20x |
| --- | --- | --- | --- | --- |
| Medium | Oil | Oil | Oil | Air |
| Brand | Olympus | Olympus | Olympus | Olympus |
| Type | UPlanSApo | PlanApo N | UPlanXApo | UPlanSApo |
| NA | 1.40 | 1.40 | 1.40 | 0.75 |
| WD | 0.13 | 0.12 | 0.13 | 0.60 |
| Coverslip | 0.17 | 0 | 0.17 | 0.17 |
| Correction collar | n | n | n | n |
| Spring | yes | yes | yes | no |
| Tube | ∞ | ∞ | ∞ | ∞ |
| Bright field | DIC | DIC | DIC | DIC |

## Lateral resolution

- λem = 515 nm

- n = 1.518

- chip pixel size = 6.45 µm

| Objective mag. | NA | Pixel size (nm) | Rayleigh limit (nm) | Nyquist pixel size (nm) | Sampled? | Sampled w/ 1.6? |
| --- | --- | --- | --- | --- | --- | --- |
| 20 | 0.75 | 323 | 419 | 182 | No | No |
| 40 | 1.3 | 161 | 242 | 105 | No | Yes |
| 100 | 1.4 | 64.5 | 224 | 97.6 | Yes | Yes |

## Axial resolution

- λem = 515 nm

- n = 1.518

| Objective mag. | NA | FWHM (nm) | Nyquist sampling (nm) |
| --- | --- | --- | --- |
| 20 | 0.75 | 1350 | 589 |
| 40 | 1.3 | 624 | 271 |
| 100 | 1.4 | 492 | 214 |

# LEDs

InsightSSI 7 colour light engine

## Spectral information

## Output power

Power (in mW) measured at the back aperture of the microscope on 06/04/2018 for four common illumination wavelengths.

# Lasers

| Wavelength (nm) | Output power (mW) | Imaging? | FRAP? |
| --- | --- | --- | --- |
| 410 | 75 | No | Yes |
| 532 | 22 | No | Yes |

# Filters

## Filter sets

| Filter set | Name | Excitation band (nm) | Emission band (nm) |
| --- | --- | --- | --- |
| Standard | DAPI | 381-399 | 411-459 |
|  | FITC | 461-489 | 501-549 |
|  | TRITC | 562-587 | 575-620 |
|  | CY5 | 621-643 | 662-696 |
| Live cell | CFP | 426-450 | 463-487 |
|  | GFP | 461-489 | 501-549 |
|  | YFP | 505-522 | 537-559 |
|  | mCherry | 562-587 | 602-647 |

## Dichroics

| Name | Channel | Reflection band (nm) | Transmission band (nm) |
| --- | --- | --- | --- |
| Quad | DAPI | 381-401 | 409-456 |
|  | FITC | 464-492 | 500-523 |
|  | TRITC | 531-556 | 564-611 |
|  | CY5 | 619-644 | 652-700 |
| Quad-mCh | DAPI | 381-401 | 409-456 |
|  | GFP | 464-492 | 500-553 |
|  | mCherry | 561-590 | 598-617 |
|  | CY5 | 625-644 | 652-700 |
| C/YFP/mCh | CFP | 400-454 | 463-487 |
|  | YFP | 496-528 | 537-550 |
|  | mCherry | 558-593 | 602-648 |

# Cameras

- Photometrics CoolSNAP HQ interline CCD

- Speed = 11 fps (full-chip), 56 fps (256x256 sub-array, 4x4 binning)

- QE = 64 %

- Read noise (r.m.s) = 4.5 (10 MHz), 5.5 (20 MHz) electrons

- Dark current = 0.001 electrons/px/s (at -30 °C)

## Field of view

- chip size = 1040 x 1392 px

- field number = 26.5 mm

| Objective mag. | Magnified chip size (µm) |  | Chip illumination area (µm) |  | FOV (µm) |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | x | y | x | y | x | y |
| 20 | 335 | 449 | 1330 | 1330 | 335 | 449 |
| 40 | 168 | 224 | 663 | 663 | 168 | 224 |
| 100 | 67.1 | 89.8 | 265 | 265 | 67.1 | 89.8 |

# Computer

- Software: softWoRx v. 5.5.0

- OS: Centos v. 5.7

- Processor: Intel(R) Xeon(R) CPU E5504 @ 2.00GHz (8 CPUs)

- Memory: 4GB RAM, 1x 500GB, 1x 2TB hard disk space

Taken from http://www.gelifesciences.co.kr/wp-content/uploads/2016/11/product-note-InsightSSI.pdf