# DV2 Spec Sheet

Contents

1	Microscope	1

2	Objectives	1

3	LED Illumination	2

4	Filters	3

5	Camera	4

6	Computer	4

# Microscope

- API Deltavision Personal DV2

- Body: Olympus IX-71

# Objectives

## Objective list

| Magnification | 100x | 60x | 40x | 20x | 10x | 4x |
| --- | --- | --- | --- | --- | --- | --- |
| Medium | Oil | Oil | Oil | Oil | Air | Air |
| Brand | Olympus | Olympus | Olympus | Olympus | Olympus | Olympus |
| Type | UPlanSApo | PlanApoN | UPlanFL N | UPlanSApo | UPlanSApo | Plan N |
| NA | 1.40 | 1.42 | 1.30 | 0.85 | 0.4 | 0.1 |
| WD | 0.13 | 0.15 | 0.20 | 0.17 | 3.1 | 18.5 |
| Coverslip | 0.17 | 0.17 | 0.17 | - | 0.17 | - |
| Correction collar | n | n | n | N | n | n |
| Spring | yes | yes | yes | No | no | no |
| Tube | ∞ | ∞ | ∞ | ∞ | ∞ | ∞ |
| Bright field | DIC | DIC | DIC | DIC | DIC | BF |

## Lateral resolution

- λem = 515 nm

- n = 1.518

- chip pixel size = 6.45 µm

| Objective mag. | NA | Pixel size (nm) | Rayleigh limit (nm) | Nyquist pixel size (nm) | Sampled? | Sampled w/ 1.6? |
| --- | --- | --- | --- | --- | --- | --- |
| 4 | 0.1 | 1610 | 3140 | 1370 | No | Yes |
| 10 | 0.4 | 645 | 785 | 341 | No | No |
| 20 | 0.85 | 323 | 370 | 161 | No | No |
| 40 | 1.3 | 161 | 242 | 105 | No | Yes |
| 60 | 1.41 | 108 | 223 | 96.9 | No | Yes |
| 100 | 1.4 | 64.5 | 224 | 97.6 | Yes | Yes |

## Axial resolution

- λem = 515 nm

- n = 1.518

| Objective mag. | NA | FWHM (nm) | Nyquist sampling (nm) |
| --- | --- | --- | --- |
| 4 | 0.1 | 91400 | 39800 |
| 10 | 0.4 | 5490 | 2390 |
| 20 | 0.85 | 1760 | 766 |
| 40 | 1.3 | 624 | 271 |
| 60 | 1.42 | 467 | 203 |
| 100 | 1.4 | 492 | 214 |

# LED Illumination

## Spectrum

Taken from imsol website 09/08/2018.

## Output power

Power (in mW) measured at the back aperture of the microscope on 06/04/2018 for four common illumination wavelengths.

# Filters

## Filter sets

| Name | Excitation band (nm) | Emission band (nm) | Eyepiece (nm) |
| --- | --- | --- | --- |
| DAPI | 382-393 | 450-500 | 450-500 |
| CFP | 418-442 | 450-500 | N/A |
| GFP | 450-490 | 500-550 | 520-550 |
| YFP | 490-510 | 520-550 | 520-550 |
| mCherry | 558-586 | 602-662 | 635/60 |
| CY5 | 633-647 | 665-705 | N/A |

## Dichroics

| Name | Channel | Reflection band (nm) | Transmission band (nm) |
| --- | --- | --- | --- |
| Quad | DAPI | 381-401 | 409-456 |
|  | FITC | 464-492 | 500-523 |
|  | TRITC | 531-556 | 564-611 |
|  | CY5 | 619-644 | 652-700 |
| GFP/mCh | GFP | 464-492 | 500-553 |
|  | mCherry | 561-590 | 598-617 |
| CFP/YFP | CFP | 400-454 | 463-487 |
|  | YFP | 496-528 | 537-550 |

# Camera

- Photometrics CoolSNAP HQ2 interline CCD

- Speed = 11 fps (full-chip), 104 fps (256x256 sub-array, 8x8 binning)

- QE = 64 %

- Read noise (r.m.s) = 4.5 (10 MHz), 5.5 (20 MHz) electrons

- Dark current = 0.001 electrons/px/s (at -30 °C)

## Field of view

- chip size = 1040 x 1392 px

- field number = 26.5 mm (22 mm for 4x)

| Objective mag. | Magnified chip size (µm) |  | Chip illumination area (µm) |  | FOV (µm) |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | x | y | x | y | x | y |
| 4 | 1680 | 2240 | 5500 | 5500 | 1680 | 2240 |
| 10 | 671 | 898 | 2650 | 2650 | 671 | 898 |
| 20 | 335 | 449 | 1330 | 1330 | 335 | 449 |
| 40 | 168 | 224 | 663 | 663 | 168 | 224 |
| 60 | 112 | 150 | 442 | 442 | 112 | 150 |
| 100 | 67.1 | 89.8 | 265 | 265 | 67.1 | 89.8 |

# Computer

- Software: softWoRx v 5.5.1

- OS: Centos v 4.8

- Processor: AMD Opteron 170 @ 2GHz (2 CPUs)

- Memory: 2GB RAM, 1x 500GB hard disk space

Calculated using eq. 16 in Wilson, T. Resolution and optical sectioning in the confocal microscope. Journal of Microscopy. 244:2 pp. 113-121 (2011)