Risk Assessment Form – RA-RTConfocal-_V1

| Title of Assessment | Room Temperature Confocal Microscope |  |  | Date of assessment |  | 25/4/17 |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Date for review |  | 25/4/18 |
| Department | Physics |  |  |  |  |  |
| Descriptions of Activities | Room temperature operation only; imaging defects in crystalline material (mainly single crystal diamond), optically detected magnetic resonance experiments |  |  |  |  |  |
| Name of those working to this assessment | Listed users in the risk assessments associated to installed lasers |  | Any others who may be affected by this assessment |  | All other users of lab G69 |  |
| Assessment carried out by | Phil Diggle |  |  |  |  |  |

| ##### Additional information | Room temperature confocal microscope with a primary excitation wavelength of 532 nm (Laser Quantum gem532, 100 mW). Individual risk assessments and safe systems of work are completed for each installed laser wavelength for this system. For optically detected magnetic resonance (ODMR) studies, a microwave source is used providing an output of X dBm. This is amplified to give a maximum output of X dBm. To deliver the microwaves to the sample, the sample is mounted onto a PCB co-planar waveguide (CPW) and a 20 um wire soldered across the top of the sample. The end of the CPW is terminated by a 50 ohm connection. A neodymium (Nd) magnet is situated above the stage and manually translated with 5-axis of freedom. The Nd magnets are cylindrical disks with a 25 mm diameter and 10 mm thickness, with a 6 mm countersunk bore in the middle. The surface field is somewhere between 1 and 2 T (2000 Gauss) – at the high end of the estimate, the 5 Gauss line situated approximately 100 mm away from the surface of the Nd magnet. |
| --- | --- |

| Foreseeable Significant Hazard | Hazard Risk | Existing control measures | Controlled residual risk | Further Action where risk remains moderate/high | By whom & when | Controlled Risk Level |
| --- | --- | --- | --- | --- | --- | --- |
| Mains electricity | Electrocution, burns. | All equipment must have an in date portable appliance testing (PAT) sticker. Wires to make connections between equipment must be electrically insulated | (Major, Unlikely) Low | none | Phil Diggle 25/4/17 | Low |
| Broken Glassware | Cuts | Samples mounted to CPW via 0.17 mm thick cover glass. Removal of sample should be done in the preparation lab to minimize any sharp glass fragments | (Minor, Likely) Low | none | Phil Diggle 25/4/17 | Low |
| Laser Light | Light affecting eye | Lasers operated with interlock connection, individual risk assessments completed. Daily use, laser light aligned and contained | (Major, Unlikely) Low | none | Phil Diggle 25/4/17 | Low |
| Microwave Radiation | Burns | ODMR only conducted by trained users named on laser RA and when properly terminated CPW mounted. SMA connector should not be installed or removed whilst microwave amp operational | (Serious, Possible) Low | none | Phil Diggle 25/4/17 | Low |
| Magnetic field | Bruising, ill-health, failure of electronic health devices | Magnetic items should not be held within the 5 Gauss line (within 100 mm of the surface of the Nd magnet) All users, and especially users with pace-makers/other electronic health devices, should not be within 100 mm of the Nd magnet | (Major, Unlikely) Low | None | Phil Diggle 25/4/17 | Low |
| Obstructed walkways | Slip, Fall, Trip | G69 lit to low intensity lighting, cables, boxes and other items must not be in the traffic flow of the lab or obstruct the base of the optical bench | (Minor, Possible) Low | none | Phil Diggle 25/4/17 | Low |
| Dirt or polluted work area | Poisoning, eye and skin irritation | Oil objective uses low autofluorescence immersion oil – contaminated lens wipes disposed of immediately to avoid others getting oil on themselves. Objective cleanup done free of gloves to avoid damaging objective. | (Minor, Possible) Low | none | Phil Diggle 25/4/17 | Low |

Work should not be carried out until the assessment is completed to a suitable & sufficient level and all required control measures are in place.

Is assessment suitable and sufficient       Yes

| Any further actions required to allow work to commence |  |  |  |  |
| --- | --- | --- | --- | --- |
| Approved By |  |  | Position |  |
| Date |  |  |  |  |

Please print a copy, sign it and keep for your records

|  | Severity of injury |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Likelihood | Superficial | Minor | Serious | Major | Extreme |  |  |  |
| Unlikely | Very low | Very low | Low | Low | Moderate |  |  |  |
| Possible | Very low | Low | Low | Moderate | High |  |  |  |
| Likely | Low | Low | Moderate | High | Very high |  |  |  |
| Very likely | Low | Moderate | High | Very high | Very high |  | Overall Risk Rating (highest level found) | Low |
| Extremely likely | Moderate | High | Very high | Very high | Very high |  |  |  |

See ‘Matrix for risk evaluation’ for further guidance.