Risk Assessment Form – RA-EC-EPR_V1.1

| Title of Assessment | Ex-Situ Electrochemical EPR Setup |  |  | Date of assessment |  | 16/11/2017 |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | Date for review |  | 05/01/2019 |
| Department | Chemistry and Physics |  |  |  |  |  |
| Descriptions of Activities | Use of the Electrochemical System coupled to Bruker EMX EPR spectrometers located in Millburn house room G78 |  |  |  |  |  |
| Name of those working to this assessment | Trained EC-EPR spectrometer users Emily Braxton, Mark Newton, Julie Macpherson, Ben Breze. |  | Any others who may be affected by this assessment |  | All other lab users |  |
| Assessment carried out by | Emily Braxton |  |  |  |  |  |

| ##### Additional information | General lab risks are detailed in the general lab risk assessment. Users are responsible for having performed a risk assessment for their chemicals used in the experiment. Bookings are restricted to trained users who are recorded in the px-epr-users web group. |
| --- | --- |

| Foreseeable Significant Hazard | Hazard Risk | Existing control measures | Controlled residual risk | Further Action where risk remains moderate/high | By whom & when | Controlled Risk Level |
| --- | --- | --- | --- | --- | --- | --- |
| Mains electricity | Electric shock and/or burns | Equipment should not be tampered with. All equipment requires an in-date PAT test. | (Major, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |
| Three-phase electricity | Electric shock and/or burns | Three-phase power supply/cables should not be tampered with. Visual inspection should be regularly carried out. | (Major, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |
| Potentiostat | Electric shock and/or burns | System should be set up correctly with no strain on the wires. Once the system is in use, avoid contact with exposed electroactive areas. | (Major, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |
| Spillage of solution | Electric shock | All chemical solutions should be kept within a plastic tray. Excess and/or unused solution should be stored/disposed of in an appropriate manner. When not in use all solutions should be kept in sealed containers. | (Major, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |
| Water leak | Electric shock and slips | Water cooling pipes/fitting should be inspected for leaks. Flooding sensors are located at all the spectrometers, which cut the water supply if a water leak is detected. Lab flooding procedure (c.f. general risk assessment) should be followed. | (Major, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |
| Microwave source | Burns to eyes/skin | Microwave power should not be on if waveguide is open. Waveguide seals should be properly fitted and replaced if damaged. Sample tubes should only be removed from the cavity when spectrometer is in Standby mode. | (Serious, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |
| Magnetic field | Bruising, ill-health, failure of electronic health devices | Magnetic items should not be held within the 5 G line. (approximately the front of the magnet yoke) All users, and especially users with pace-makers/other electronic health devices, should not cross the 5 G line. (approximately the front of the magnet yoke) | (Major, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |
| Broken glassware | Cuts and/or poisoning | Take care, dispose of sharps and broken glass in the approved receptacles provided using appropriate PPE. | (Minor, Likely) Low | None | Emily Braxton 16/11/2017 | Low |
| Obstructed walkways | Slip, fall | Do not block walkways and be aware of risk | (Minor, Possible) Low | None | Emily Braxton 16/11/2017 | Low |
| Dirt or polluted work area | Poisoning, eye and skin irritation | Keep work areas clean and tidy | (Minor, Possible) Low | None | Emily Braxton 16/11/2017 | Low |
| Exposure to mercury | Poisoning | Handle reference electrode carefully to prevent the release of mercury. | (Major, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |
| Over pressuring of system by formation of gaseous by-products | Explosions and/or injury | Safe blow off points to be incorporated into the system to prevent explosive over pressure being reached | (Major, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |
| Use of gases for degassing | Asphyxiation | Only use gases in a well ventilated region. Lab is fitted with oxygen depletion sensors. | (Major, Unlikely) Low | None | Emily Braxton 16/11/2017 | Low |

Work should not be carried out until the assessment is completed to a suitable & sufficient level and all required control measures are in place.

Is assessment suitable and sufficient       Yes

| Any further actions required to allow work to commence |  | None |  |  |
| --- | --- | --- | --- | --- |
| Approved By | Dr Ben Breeze |  | Position | RTP - Manager |
| Date | 17/11/17 |  |  |  |

Please print a copy, sign it and keep for your records

|  | Severity of injury |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Likelihood | Superficial | Minor | Serious | Major | Extreme |  |  |  |
| Unlikely | Very low | Very low | Low | Low | Moderate |  |  |  |
| Possible | Very low | Low | Low | Moderate | High |  |  |  |
| Likely | Low | Low | Moderate | High | Very high |  |  |  |
| Very likely | Low | Moderate | High | Very high | Very high |  | Overall Risk Rating (highest level found) | Low |
| Extremely likely | Moderate | High | Very high | Very high | Very high |  |  |  |

See ‘Matrix for risk evaluation’ for further guidance.