## VT Calibration
Temperature calibration chart of 4mm HX High Temperature probe (H160530)
Calibration Date | High Temperature
2017-05-16 | [5 empty] | by Kenneth Inglis (FB Group, Liverpool)
| MAS 10 kHz | Fully packed, with a zirconia endcap, equilibrium time > 10 min
Bearing gas is VT gas. No extra VT gas.
N2 gas for bearing. Purge air for cooling
Show shim coil temperature with "coiltemp" command
KEEP SHIM LOWER THAN 80 C
Ref Beckmann & Dybowski JMR 2000 379 using δ(iso) = -3490.4 ppm at 4 kHz
Pb(NO3)2
gas (l/h)/ heater (%) | Sensor T / K | Peak shift / ppm | Real T / K | Real T / °C | T gradient / ppm | T gradient / K | Shim T / C | Comment
135, 0 | 293.6 | -3484.6 | 302.5066 | 29.5066 | 3 | 3.9578 | 27 | Peak referenced at 4 kHz
135, 5 | 323.2 | -3461.4 | 333.1135 | 60.1135 | 3.3 | 4.3536 | 30
135, 12 | 373.2 | -3422.1 | 384.9604 | 111.9604 | 4.6 | 6.0686 | 33
135, 20 | 423.2 | -3383 | 436.5435 | 163.5435 | 7.6 | 10.0264 | 44
135, 26 | 473.2 | -3342.9 | 489.4459 | 216.4459 | 11 | 14.5119 | 53
135, 32 | 503.2 | -3319.7 | 520.0528 | 247.0528 | 12.6 | 16.6227 | 60
79Br on KBr
Sensor T /oC | Sensor T /oC | Sample T/K | Sample T/oC
573 | 300 | 557 | 284
Tsample = 1.04*Tsensor -2.73
[9 empty] | 293.6 | 302.5066
[9 empty] | 323.2 | 333.1135
[9 empty] | 373.2 | 384.9604
[9 empty] | 423.2 | 436.5435
[9 empty] | 473.2 | 489.4459
[9 empty] | 503.2 | 520.0528

## MAS speed vs. Sample T
MAS | Shift /ppm | Sample T /K | Sample T /C | Difference
4 | -3490.4 | 294.8549 | 21.8549 | 0
8 | -3487.1 | 299.2084 | 26.2084 | 4.3536
10 | -3484.6 | 302.5066 | 29.5066 | 7.6517
14 | -3476.8 | 312.7968 | 39.7968 | 17.942