Safety Notes for users of the UK 850 MHz Solid State NMR Facility

These notes contain safety information for visitors to the UK High Field Solid-State NMR National Facility at the University of Warwick

More Health and Safety information can be obtained from the University of Warwick, Health, Safety and Wellbeing webpages https://warwick.ac.uk/services/healthsafetywellbeing/ and in particular.

https://warwick.ac.uk/services/healthsafetywellbeing/guidance/handspolicy.  It is suggested that new visitors familiarise themselves with these pages.

These notes have been prepared by the local management team of the UK High Field Solid-State NMR Facility.

Nuclear Magnetic Resonance Laboratory Safety Advice

For Research Visitors to the UK High Field Solid-State NMR Facility

This advice supplements other University and Department documents concerning safety and good practice in Physics laboratories.  It is intended to indicate potential hazards associated with the Nuclear Magnetic Resonance (NMR) laboratories in the Physics Department and not to replace any safety documents.

Although you are here to work at the National Facility you need to be aware of the other NMR instruments at Millburn House.  Listed are the main hazards of working in the Nuclear Magnetic Resonance laboratories.  You should be aware or seek advice on any additional hazards associated with your work including the samples you are working with.  Do not leave your common sense at home and remember that this is not intended to be an exhaustive list.  The NMR laboratory is a safe environment.  We are a large group of experienced postgraduates and staff.  When seeking any advice feel free to ask anyone in the group.

IF IN DOUBT ASK.

1/ Magnetic Fields

Although most of the magnetic field of a superconducting magnet is constrained within the magnet Dewar they do have a large stray magnetic field.  There is no medical evidence that a large static magnetic field is harmful.  However, if you are in the NMR laboratories assume you are in a magnetic field greater than 10 gauss and potentially much greater.  It is worth noting that many magnets in the main hall are unshielded magnets with a very large stray field.  If you have a heart pacemaker, insulin pump, other medical implant or metal implants (orthopaedic plate etc) you will need to inform the local management of the national facility.

The magnetic field at the bottom base plate of all NMR magnets is of the order of 0.3 Tesla.  Any ferromagnetic object (keys, spanners) will experience a very strong attractive force if bought close to close to a magnet.  Large items such as oscilloscopes and RF amplifiers will also be pulled towards the magnet if they are positioned too close.  Do not bring any object that is likely to be attracted towards a magnet close to the magnet.  Personal items can be left in the pod (850 MHz) and lobby area (1 MHz).

DO NOT TAKE ANY EQUIPMENT OR METALLIC OBJECTS THAT WILL BE ATTRACTED TO THE SUPERCONDUCTING MAGNET CLOSE TO THE MAGNET.

2/ Radio Frequency Radiation

While electromagnetic radiation has the potential to cause local heating of the body, there is sufficient shielding of the equipment and cables to screen you from this electromagnetic radiation but check all the connections before carrying out an experiment.

MAKE SURE ALL PROBE, AMPLIFIER AND CABLE RF CONNECTIONS ARE MADE CORRECTLY.

3/ Cryogenic Hazards

Superconducting magnets contain both liquid nitrogen (at 77K) and liquid helium (at 4.2K).  Refilling a magnet with either cryogen may take place without the need to stop an experiment.  If you are present during a refill and need to go near the magnet, take advice from those refilling the magnet.

Although rare there is the possibility of the superconducting magnet windings going normal (resistive)  this is referred to as a quench.  This can happen spontaneously but is usually associated with a cryogen refill or the magnet being physically disturbed.  If a quench occurs liquid helium is immediately boiled off.  As the helium gas warms up it expands, this expanding gas will displace air in the room causing a reduction in oxygen levels.  The helium gas will be cool enough to condense water vapour in the atmosphere causing a fog. This will make it difficult to see.  A quench is obvious  there will be the sound of a large volume of gas escaping and fog appearing above the magnet.  In addition, there is the possibility of nitrogen gas also being released, thus adding to the reduction of oxygen levels.

Oxygen levels in the laboratory are permanently monitored.  The low-level oxygen alarm has two stages, 1) a low whooping sound for slightly below normal levels and 2) a high whooping sound for oxygen levels that are dangerously low.  For both stages amber beacon lights in the laboratory will flash on.  A panel with the oxygen concentrations is located outside the main hall - you will be shown the location of this panel.

IF EITHER ALARM SOUNDS LEAVE THE LABORATORY IMMEDIATELY.

It is not expected that facility visitors will need to work with liquid cryogens. If you do need to use liquid cryogens, please talk to the local management and study the risk assessments and other associated documents for handling liquid helium and nitrogen available on the national facility website.

4/ Compressed gases

Magic angle spinning uses compressed gases (air or nitrogen) at a pressure of up to 8 bar.  Check all gas connections to probes are made properly.  Before disconnecting gas lines to probes check that spinning has stopped and that control regulators are turned to zero pressure.  After first consulting the facility manager, it may rarely be necessary to spin a rotor in a probe outside of the magnet.  It is possible for a rotor to jump out of the stator or to break.  If spinning a rotor outside of the magnet wear eye protection and make sure any openings in the probe body are closed.

5/ Chemical Hazard

Check how to safely handle your samples and take appropriate precautions. Users are reminded that it is a condition of use of the national facility facility that a COSHH (“Control of substances hazardous to health”) form must be submitted (by email to the facility manager) at least one week before the allocated experimental time.  COSHH sheets can be consulted for “brought in” chemicals used, e.g., as setup samples.  It is important that all rotors, sample containers etc. are clean when you have finished.  Other users will not know how toxic a powder is if it is left in a rotor and will have to treat it as a toxic and dangerous sample when disposing of it.

REMEMBER TO CLEAN UP AFTER YOURSELF.

6/ Trip Hazards

There must be cabling and compressed gas lines running on the floor for all spectrometers.  Always watch out for this and do not stand on any cable or gas line.  There may also be other small items such as frequency counters and spinning speed controllers on the floor.  Note that the 850 MHz magnet sits on a raised plinth.

7/ Rotor Packing

If not handled correctly, it is possible that MAS rotors can break during packing, potentially causing injury. However, damage or injury is very unlikely if the correct packing tools are correctly used. If in doubt, seek advice.

8/ Fire Exits

Millburn House is a complex building, and you will need to familiarise yourself with the FIRE EXITS.  The fire assembly point is located at the front of the building close to Millburn Hill Road.

9/ Fire, Accidents and near misses

Should all be reported (see later pages for first aiders).  Serious accidents and incidents when no one is immediately available should be reported to University Security using the emergency telephone number int. 22222. Ext. 024765 22222

10/ Magnetic Resonance laboratory address

Magnetic Resonance Centre

Millburn House

Millburn Hill Road

University of Warwick Science Park

CV4 7HS

NB: Post should NOT be sent to the laboratory address, but rather to:

Department of Physics

University of Warwick

Coventry CV4 7AL

Working out of Normal Hours

The Department has installed an automatic security system (Salto). All perimeter access doors are locked at all hours of the day.  Access to the national facility is also access controlled.

The system also logs arrivals and departures.

First Aid At Millburn

A first aid box is kept in G71 , the preparation laboratory.  If you require a first aider either find a member of the local management or call community safety, int. 22222. Ext. 024765 22222.