Laser Risk Assessment

- LRA1   Background Information

| Background information |  | Date 28/08/2019 Assessors M.J. Coak, P.A. Goddard |
| --- | --- | --- |
| Describe the product and application | 532 nm visible CW laser mounted in closed enclosure along with spectrometer. Laser light is coupled in and out of enclosure through an optical fibre. This fibre is mated or held into a pressure cell or cryostat to measure the fluorescence of a ruby pressure gauge. The overall system is Class 2 during normal use, as all high intensity beam sections are enclosed and interlocked. If the bolted enclosure is opened and the built in interlocks defeated, the system becomes Class 3B |  |
| Describe the laser | 532 nm (visible, green) continuous wave diode laser, 40 mW at supplied (maximum) power. ZLaser ZM18 series with attached heatsink mount. |  |
| Describe the beam delivery system | Within enclosure, beam directed onto a steerable mirror then coupled into one end of a bifurcated optical fibre by focussing optics. Beam travels down the fibre, out of the enclosure to externally attached pressure cell or cryostat, then back up the fibre and into the enclosure, where the third end of the fibre is mated into a spectrometer unit (Ocean Optics HR4000). Where the beam emerges from the enclosure, its maximally achievable power is less than 1 mW at the fibre end. This drops to 0.1 mW at 10cm as the beam is strongly divergent. Electrical faults or misalignment cannot increase this – this is the highest output power possible. |  |
| Describe the laser process | When the fibre has been held into a pressure cell, the laser is switched on and a fluorescence spectrum collected (10 seconds). An XY mount allows the fibre to be translated with respect to the cell while collecting in order to optimise the signal while keeping the fibre securely clamped in place. The laser is then switched off and the cell removed. The fibre can also be mated to the upper flange of a cryostat in a similar manner and procedure. Two accompanying Safe System of Work documents cover the procedures for this normal use scenario, and periodic realignment or adjustment procedures. |  |
| Describe the environment | For use primarily in the high magnetic field lab of the Superconductivity and Magnetism group, and in the central x-ray facilities. May also conduct work inside dedicated laser labs. Rooms are well lit. Realignment operations (where the system becomes Class 3b) carried out in keycode access controlled labs with windows masked off with blinds. Signs forbidding entry are to be displayed on the door during realignment operations. |  |
| Who uses the product or who could affect its operation | Normal use, maintenance: Trained users authorised by P. Goddard only. Users must be separately and specifically trained and authorised for the realignment/maintenance operations, which require operating the laser with the enclosure open and interlocks overridden. |  |
| Identify the parts of the life cycle relevant to the risk assessment | Normal operation, Maintenance |  |

- LRA2    Risk Assessment

| Assessment Number: 1 Assessment Date: 2019-08-28 Assessed by: Review Date: Continuous |  |  |  | Activity/Facility Assessed: Ruby spectrometer laser setup |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| STEP 1 | STEP 2 | STEP 3 |  |  |  |  |
| List of significant hazards | List groups of people who are at risk | Describe existing controls | Are these controls adequate ? |  | What is the residual risk factor from these hazards? | Actions Required (See separate form LRA3) |
| The Laser: risk of retinal damage | Users – trained and designated personnel only | The laser itself is a Class 3B, CW, visible. This rating only applies during realignment operations which require the enclosure to be opened and interlocks overridden. These operations are covered in accompanying safe system of work documents. Beam is kept horizontal by robustly mounted optical components. Enclosure terminates beam and prevents insertion of eye into beam area (lid on top of the box enclosure opens). Appropriate laser safety goggles provided for these procedures. During normal use, overall system is a Class 2 only, see below. | Yes |  | Low | - |
| Beam Delivery: risk of retinal damage | All people in lab environment during use (10 second measurements, with oversight. Not left running unattended) | As the laser and hazardous beam sections are contained in an interlocked enclosure, the overall system is a Class 2 laser system. This then does not require any additional safety measures beyond avoiding prolonged direct eye exposure. A safe system of work document covers this normal use scenario. Clear warning signs are displayed advising of the laser hazard and to avoid staring into the beam. | Yes |  | V.Low | - |
| The Laser Process: risk of retinal damage | As above | As above | Yes |  | V.Low | - |
| Environment & People: risk of retinal damage | As above | As above | Yes |  | V.Low | - |

| Examples of hazards |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| People | Noise |  | Fumes | Explosion | Hot objects | Confined spaces | Optical components failure |
| Poor instruction, training | Fire |  | Moving parts of machinery | Radiation hazards | Ozone | Poor lighting | High vacuum systems |
| Inadequate procedures | Lifting of heavy objects |  | Pressure systems | Electromagnetic radiation / X-rays | Trailing leads | Chemicals | Dust |
| Unsuitable equipment | Electricity |  | Falling objects | Inadequate optics mounts | Inadequate coolant pipe connections | Biohazards | Slip, trip hazards |

- LRA3    Action tracking.

| Significant Hazards Identified (from form LRA2) | STEP 4 Actions Required | Date for Action | Completed by (name / date) |
| --- | --- | --- | --- |
| The Laser: | - | - | - |
| Beam Delivery: | Initial commissioning and testing of the laser system, to be carried out under the supervision of the departmental laser safety advisor, will verify the safety class of the overall system. Safety procedures and controls will be reviewed and the output beam power through the optical fibre tested and measured. Inspection of the system by the Laser Protection Adviser, and review of all safety documentation by University Safety Office. | Completed Completed | M Coak, P Goddard, J. Llyod-Hughes, July 2019 P Goddard, J Brannon, Laser Protection Advisor, Aug 2019 |
| The Laser Process: | - | - | - |
| Environment & People: | - | - | - |