Risk Assessments for Equipment – Guidance

Risk assessments must be completed for each major piece of equipment. Based on the completed assessment, any additional training courses required to use the equipment, beyond those already defined for accessing the space, as well as requirements for competency signoff should be noted herein. Please consult the training matrix spreadsheet for courses. The main outcome of an equipment risk assessment should be a standard operating procedure that must be used for normal operation of the equipment. At a minimum, the equipment risk assessment must include the emergency switch off procedures and include any specific PPE required.

Where appropriate, consider if there is a need for “expert users”, such as those who could maintain or use the equipment beyond the stated standard operating regime. If so it should be clearly stated who these users are and what additional operational rules/training are needed.

A physical hard copy of the standard operating procedures and any approved user list should be kept with the equipment so that it is accessible to all users with the electronic copy stored on Quemis under the room in which it is located.

## Risk Assessment Form

| Title of Assessment | Risk Assessment for High-Pressure Image Furnace Upgrade |  | Date of assessment |  | 19/04/2023 |
| --- | --- | --- | --- | --- | --- |
| Department | Physics |  | This risk assessment must be reviewed whenever a significant change or an accident occurs |  |  |
| Descriptions of Activities | Crystal growth at high pressuring using a modified commercial image furnace. |  |  |  |  |
| Name of those working to this assessment | Tom Orton Daniel Mayoh Geetha Balakrishnan | Others who may be affected by this assessment |  | Anyone working in the lab (P124) at the time of use. |  |
| Assessment carried out by | Tom Orton |  |  |  |  |
| Additional information | The space risk assessment and image furnace risk assessment/SOP for P124 must also be read in conjunction with this document. |  |  |  |  |

## Standard Operating Procedures

Will be written once equipment is out of development phase and put into use by trained personnel.

## Emergency Switch Off Procedure

Leave room, if safe to do so, switch the main circuit breaker off. There is no need to interfere with the equipment any further.

## Additional Training Requirements

Users must be trained to use this equipment in addition to the training is required to enter and work in P124.

| Hazards and how they may cause harm | Who may be at Risk? | Existing Control Measures | Current Risk Level (VL,L,M,H,VH) | Where current risk is M, H or VH, what additional Control Measures are required? | Action required by whom & by when? | Final Risk Level |
| --- | --- | --- | --- | --- | --- | --- |
| Exploding quartz tube - can cause cuts, burns and temporary loss of hearing. | Everyone in room. | All tubes are rated at least 1.5 times their working pressure. All tubes are tested in-situ to their working pressure while surrounded by a blast-shield before they are used. Each tube has a number which can be looked up in the manual to see what the maximum working pressure of the tube is. During normal operation the whole setup is double encapsulated in protective screening to stop broken pieces of glass escaping. The furnace encapsulation/protective screening is all interlocked. | Low | None. | Users must check the setup for defects before starting. The users should also check that there is no visible damage to the equipment before starting. | Low |
| Gas Regulator failure might cause sudden release of gas into the surrounding environment – can cause cuts, aneurisms and temporary loss of hearing. | User/people nearby. | Gas Regulators are thoroughly checked for compliance, in-line with local ruling. High pressure connection to cylinders is only to be made by competent technical staff. | Very Low | None. | Technical staff must check all regulator connections and correctly leak-test them before the equipment is handed to the user(s). | Very low |
| High pressure hose fails – can cause whipping/flailing of hose which may cause injury. | People very close by (<1 m). | Hoses are all braided and rated to 3 times their working pressure (15 times their burst pressure). Hoses are installed and used as per the manufacturer’s instructions. | Very low | None. | Periodic inspection by competent persons | Very Low |
| Fire from hot pieces of exploding tube – if the tube fails when the furnace is in operation it can send very hot pieces of glass hurtling across the room ready to burn, singe and cut those who accidentally are in its path. | Everyone in room. | Furnace is manufactured with blast screening in place to stop anything from being ejected during use. It is still possible for hot pieces of glass to fall to the floor underneath the furnace. | Low | None. | Tubes should be tested using blast screen before they are used for a crystal growth. | Low. |
| High Pressure (HP) gas entering lower pressure part of system (including regulator failure). – This would result in a large bang or whoosh that could startle and cause distress to some people. | Users/operators of the furnace. | All parts of the gas management system are protected by over-pressure (safety) valves. All of the valves exhaust though high-pressure tubing to outside the building, so no release of gas into the room can happen in this condition. | Very low | None | Safety valves should be periodically tested/inspected. | Very low |
| Incorrect assembly causing escape of HP gas. – Escape of hot gas at high velocity inside furnace can cause rapid seal failure which could also result in fire or burns. | Users/other people in room. | Instruction manual and user training is mandatory for all people wanting to use this equipment. The setup must be leak-checked when cold to ensure that there are no leaks. | Very Low | None. | Users must check for leak-tightness every time before starting. | Very low. |
| Explosion in furnace. – This would cause a loud band and send bits of broken glass across the room. The glass might also be very hot. This can cause burns and cuts. | Users/other people in room. | All furnace functions are tested before beginning. All gas handling parts are checked for leak-tightness before beginning. The furnace is double screened and interlocked. | Low | None. | Users must setup the furnace as per the instruction manual and check for leak-tightness before starting. | Very low |
| Leak of gas into room – gas from the lab supply or cylinder can leak into the room. This may cause oxygen depletion. | Anyone in the room. | Lab is very large in volume in relation to the amount of gas that can be leaked into the room. In a catastrophic failure scenario the O2 level in the room will not drop below ##see O2 depletion calculation from previous assessment## | Low | None. | Gas equipment is inspected periodically. | Low |

Work should not be carried out until the assessment is completed to a suitable & sufficient level and all required control measures are in place.

| Overall Final Risk Rating (Highest level in final column above) |  | Low |  |
| --- | --- | --- | --- |
| Additional Comments from Risk Assessor (e.g. funding or practical implications) |  |  |  |  |  |
| Any further actions required to allow work to commence |  |  |  |  |  |
| Approved By |  |  |  | Position |  |
| Date |  |  |  |  |  |

Please print a copy, sign it and keep for your records

|  | Severity |  |  |  |  |  |  | Risk Level |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Likelihood | Superficial | Minor | Serious | Major | Extreme |  | Very low | Acceptable risk - no action required |
| Unlikely | Very low | Very low | Low | Low | Moderate |  | Low | Tolerable risk - further control measures not required, but status must be monitored |
| Possible | Very low | Low | Low | Moderate | High |  | Moderate | Further control measures required to reduce risk as far as is reasonably practical |
| Likely | Low | Low | Moderate | High | Very high |  | High | Urgent action required to allow activity to continue |
| Very likely | Low | Moderate | High | Very high | Very high |  | Very high | Risk intolerable - activity must cease until the risk has been reduced |
| Extremely likely | Moderate | High | Very high | Very high | Very high |  |  |  |

See ‘Matrix for risk evaluation’ for further guidance.