Gas Admittance to Vacuum Systems

High pressure gas is a hazard! Do not use any gas bottle unless you have been trained in gas admittance and how to use the regulators.

If you suspect there is a problem with a regulator or bottle, contact Gavin Bell, Alan Burton or another member of technical staff. Do NOT attempt to fit or fiddle with regulators on large bottles. You MAY change regulators on small bottles once trained (bottles screw in an out of a needle valve).

High pressure gas is stepped down in pressure from 12 bar (small bottles) or up to 200 bar (large bottles) so that we can admit gas at milli bar pressures or less. (1 bar = atmospheric pressure)

Stage 1: regulator

Regulators allow us to step down from the high pressure in the cylinder to a pressure around 1 bar. They also protect us against over-pressure and are essential pieces of safety equipment. Regulators must be replaced every few years. The dial on a regulator is not a valve: normally clockwise motion increases the pressure in the regulator.

Stage 2: gas line

The vacuum systems have “gas lines” which are high vacuum compatible sections which we fill with a reservoir of gas via the regulator. The pressure in the gas line is not usually measured directly but typically will be some fraction of a bar (still far higher than the chamber pressure!). Details of following procedure depend on the vacuum system.

- We never open the bottle to the gas line directly. We use the small volume of gas in the regulator itself to fill the gas line to a fraction of a bar pressure

  - Open regulator to bottle, adjust pressure to a few bar.

  - Close bottle off from regulator.

  - Open valve from regulator to gas line.

  - Pressure in gas line = pressure in regulator

- Pump and flush gas line: one or two cycles of rough pumping and flushing the line with gas from the regulator admitted as above.

- Fill gas line: normally two or three “fills” as above.

Stage 3: the chamber

To go from the gas line (fraction of a bar) to chamber (around 10-9 bar) we use a leak valve. These have a fine thread which when turned separates two metal surfaces in the valve, admitting a very small pressure. Details of following procedure depend on the vacuum system.

- Open the leak valve by the known number of turns before it starts to actually open.

- Monitor the pressure in the chamber as you continue to open the leak valve.

- Adjust the leak valve for the desired pressure noting there may be a time lag.

Never over-tighten a leak valve when closing up! The threads are fine and can be wrecked by clumsy over-tightening.