Risk Assessment Form

| Title of Risk Assessment | Handling of cryogenic fluids |  |  | Date of assessment |  | 26/02/19 |
| --- | --- | --- | --- | --- | --- | --- |
| Department | Physics Department |  |  | Date review due |  |  |
| Description of Task/Process | The filling, dispensing and use of cryogenic fluids |  |  |  |  |  |
| Assessment carried out by | Surface Science Group |  |  |  |  |  |

| ##### Additional information | This is a risk assessment for Dealing with a small amount of liquid nitrogen (roughly 5 litres at one time) used for Hall measurements in Room 423A. We will not be dealing with pressurised containers or liquid helium. |
| --- | --- |

| Hazards and how they may cause harm | Who may be at Risk? | Existing Control Measures | Current Risk Level (VL,L,M,H,VH) | Where current risk is M, H or VH, what additional Control Measures are required? | Action required by whom & by when? | Final Risk Level |
| --- | --- | --- | --- | --- | --- | --- |
| Cold burns to the body | User, bystander | - PPE - cryogenic gloves must be worn. BS EN 420, BS EN 511, BS EN 388 - Enclosed footwear must be worn. If boots are worn, trousers should be worn outside the boots to prevent liquid running into the top of the boot | M | - Training to be carried out by competent person. In case of spilling on to shoes, take them off so LN2 cannot be held against the skin. | - The lab supervisor to ensure that users are trained prior to permitting them to handle cryogens. | Low |
| Cryogenic splash in eyes (filling and use) | User, bystander | - Cryogenics handled in restricted area or an area which is located away from other activities. - PPE – eye protection to be worn BS EN 166 grade 3, goggles or full-face visor - Training to highlight this need. - Signage. | L |  |  | Low |
| Asphyxiation when filling | User, bystander | - Labs suitably designed with adequate ventilation. - Remain attentive to filling operation and do not overfill. - We use very small amounts of liquid nitrogen (around 5 litres) | L | - In case of liquid nitrogen being spent on the floor of P423A or P424, Leave the room for adequate time and put NO ENTRY sign on the door | - user | Low |
| Cryogenic burns from contact with cold surfaces | User | PPE - cryogenic gloves to BS EN 420, BS EN 511, BS EN 388. Training to alert user to this hazard. | L |  |  | Low |
| Injury to back, arms, legs when moving trolley (slips, trips) | User, bystander | - Dewars have wheels that can aid transport from one location to another. - Manual handling training. Awareness of push rather than pull. Awareness of uneven surfaces likely to cause topple. - Route to be pre-planned, to avoid slopes and stairs. | M | Take care wheeling medium sized dewar out of the goods lift. Store dewars out of the central part of P423A – keep in recess between glove box and sink. | Estates Project Managers and Department involved in design of facility. | Low |
| Asphyxiation when using or transporting | User, bystander | - Work in well-ventilated area. If working in an enclosed space use oxygen depletion alarm. - When transporting a dewar in the goods lift. it must be transported separately and place a sign to prevent others from using lift at the same time. | M | Persons to be trained in emergency procedures if alarm sounds. Our medium sized wheeled must go in the goods lift – ensure no-one travels in the lift with the dewar (use signage on dewar and summon lift from level 4). | Local manager / supervisor to ensure persons are trained in the emergency procedures prior to any work taking place involving cryogens. | Low |
| Obstruction of walkways – fire escape issues | User, bystander | Users are aware that they need to maintain thoroughfares at all times. | L | Regular checks of thoroughfares to be carried out by departmental staff. To report issues as they arise. | Any person in the Department (at all times) | Low |

Work should not be carried out until the assessment is completed and all required control measures are in place.

| Overall Final Risk Rating (Highest level in final column above) | M |
| --- | --- |

| Additional Comments from Risk Assessor (e.g. funding or practical implications) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Approved By | G Bell |  | Position | Reader |
| Date | 26/02/2019 |  |  |  |

Please print a copy, sign it and keep for your records

|  | Severity |  |  |  |  |  |  | Risk Level |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Likelihood | Superficial | Minor | Serious | Major | Extreme |  | Very low | Acceptable risk - no action required |
| Unlikely | Very low | Very low | Low | Low | Moderate |  | Low | Tolerable risk - further control measures not required, but status must be monitored |
| Possible | Very low | Low | Low | Moderate | High |  | Moderate | Further control measures required to reduce risk as far as is reasonably practical |
| Likely | Low | Low | Moderate | High | Very high |  | High | Urgent action required to allow activity to continue |
| Very likely | Low | Moderate | High | Very high | Very high |  | Very high | Risk intolerable - activity must cease until the risk has been reduced |
| Extremely likely | Moderate | High | Very high | Very high | Very high |  |  |  |

See ‘Matrix for risk evaluation’ for further guidance.