## Slide 1: Particle Identification Warwick Week
Antonis Papanestis
RAL – STFC
(with borrowed material from S. Easo)

## Slide 2: Outline
Introduction
Particle ID techniques:
  Time of flight
  Transition radiation
  Cherenkov radiation
    RICH detectors
    DIRC detectors
  dE/dx
Give examples of experiments and detectors

## Slide 3: Introduction (i)
A “generic” particle physics experiment has “built-in” a lot of particle ID for:
Electrons, photons, neutrons and muons

## Slide 4: Introduction (ii)
Particle ID for hadrons
  Distinguish between kaons, pions, protons, deuterons
    Required for flavour physics
Specific experiment requirements
  Distinguish between electrons and charged pions
  Distinguish between electrons and muons
    For neutrino experiments
    For rare decay searches
before and after RICH PID

## Slide 5: Time of flight
Time difference for two particles with masses m1 and m2 for length L
Mass resolution depends on time resolution of detectors and distance L
Typical values: L=3.5 m, Dt=100 ps, for 3 s separation, Pmax=2.1 GeV/c
Requires fast detectors: scintillation, Cherenkov, fast collection of ionisation

## Slide 6: PID with TOF
Expected PID performance for various path lengths and time resolutions
TORCH, a proposed future detector, can push p/K separation to 9 GeV/c with L~10 m and time resolution per track ~ 20 ps

## Slide 7: TOF: ALICE experiment
Resistive plates made of glass.
2 X5 gaps : 250 mm
Readout by HPTDC chip (High Performance Time to Digital Converter)
s=50.8ps
3σ π/K separation up to 2.2 GeV/c and K/p separation up to 4 GeV/c.
http://aliceinfo.cern.ch/Public/en/Chapter2/Chap2_TOF.html

## Slide 8: Transition Radiation (TR)
Transition Radiation: Radiation in the x-ray region when ultra
relativistic particles cross the boundary
between 2 media with different dielectric constants.
Mainly for e-p separation in 0.5 GeV/c  200 GeV/c.
Full explanations and the derivations from Maxwell’s equations:
NIMA 326 (1993) 434-469 and references there in.
f
q
The radiation is peaked at a small angle:
q = 1/g
g
wi =plasma freq of medium i

## Slide 9: TR Detection
Integrating the previous equation, one can get , for xg=0,
g = E/m of the particle. This makes PID possible by measuring W.
Lighter particle give larger signal
wf = plasma frequency = 28.8 (r Z/A)0.5 eV
  r=density, Z=atomic weight, A=atomic number
For example, for wf = 0.02 keV and g = 5000, most of the photon energy is
in the range 10 keV< w <100 keV (ie. 0.1 wc< w < wc
where w c = cut-off frequency ).
Number of photons produced =
For wc=100 keV and w=1 keV, N= 0.03 for a single surface.
Hence to get sufficient number of photons , large number of interfaces are used:
a stack of many foils with gaps in between.

## Slide 10: TRD Example 1: HELIOS experiment (NA34)
The minimum thickness of the foils and air gaps are determined by the size of the
‘formation zone’ and the interference effects.
( typically foils can be 10-20 mm thick and are made of polypropelene)
Behind a TRD foil stack there is a MWPC or drift chamber where the
TRD signal is detected along with the signal from the charged track.
Drift space = 10 mm,
Anode space = 6 mm
Drift time=0.51 ms
May use FADC or discriminators

## Slide 11: TRD Example 2: ATLAS TRT
Barrel and endcap TRT
BdJ/y Ks
In an average event, energy deposit from:
Ionization loss of charged particles ~ 2.5 keV
TR photon > 5 keV.
(Photon emission spectrum peaks at 10-30keV)
Blue dots: ionizing hits
Red dots : TR hits
TRT straw wall design

## Slide 12: Cherenkov radiation

## Slide 13: Basics of Cherenkov radiation
q
Charged particle
photon
Refractive index
b= velocity of the charged particle in units of speed of light (c) vacuum
P, E ,m = momentum, energy, mass of the charged particle.
CM = speed of light in the medium (phase velocity) ,
Eph = photon energy, l=Photon wavelength.
Theory of Cherenkov Radiation: Classical Electrodynamics by J. D. Jackson (Section 13.5)

## Slide 14: History of Cherenkov radiation
The formula cos (q ) =1/(nb) was already predicted by Heaviside in 1888
~1900: ‘Blue glow’ seen in fluids containing concentrated Radium (Marie & Pierre Curie)
Pavel Alexeevich Cherenkov (1904-1990): Lebedev Physical Institute of the Russian Academy of Sciences.
Discovery and Validation of Cherenkov Effect : 1934-37
Full Explanation using Maxwell’s equations: I.M. Frank and I.E. Tamm in 1937
Nobel Prize in 1958: Cherenkov, Frank and Tamm.

## Slide 15: First experiments
1: vessel with liquid
2: mirror
3: Cherenkov photons towards the photographic plate
Typical Apparatus used by Cherenkov to study the angular distribution of Cherenkov photons.
(Incident g ray produces electrons by Compton scattering in the liquid).
P. Cherenkov established that:
Light Intensity is proportional to the electron path length in the medium.
Light comes only from the ‘fast’ electrons above a velocity threshold, in his Apparatus.
Light emission is prompt and the light is polarized.
The wavelength spectrum of the light produced is continuous. No special spectral lines.
The angular distribution of the radiation, its intensity, wavelength spectrum and its dependence on the refractive index agree with the theory proposed by his colleagues Frank and Tamm.

## Slide 16: Refractive index
Glass
Air
Dependence of the refractive index on photon frequency means the Cherenkov angle is not constant
The spread of angles is referred to as Chromatic Error

## Slide 17: Number of photons
Number of emitted photons:
The number of photons increases with energy
Between two wavelengths:
Using frequency instead of wavelength
Typical example: Charged particle with momentum of few GeV/c or more emitting Cherenkov photons with few eV of energy

## Slide 18: Detected photons
For lL=400 nm and lH=700 nm
If the photons are reflected by a mirror with reflectivity R(l), and detected by a photon detector with quantum efficiency QE(l) through a window with transmission T(l)
And if we assume the mean angle qc
N0 is a figure of merit for a Cherenkov detector; e.g. N0=200/cm is a good value

## Slide 19: A variety of Cherenkov detectors
Detector design:
  Threshold Counters
  Imaging detectors
    Differential Cherenkov detectors
    Ring Imaging Cherenkov detectors (RICH)
    Detector for Internally Reflected light (DIRC)
Types of photon detectors:
  Gas based
  Vacuum
  Solid state
Applications:
  Accelerator HEP detectors
  Astroparticle Physics detectors
  Neutrino detectors

## Slide 20
Threhold Cherenkov Counters
Signal produced from only those particles which are above Cherenkov Threshold.
Basic version: Yes/No decision on the existence of the particle type.
One counts the number of photoelectrons detected.
Improved version: Use the number of observed photoelectrons or a calibrated pulse height
to discriminate between particle types.
For typical detectors: No = 90 cm-1,
Nph per unit length of the radiator = No * (m12 – m22)/(p2+m12)
At p= 1 GeV/c, Nph per unit length = 16 /cm for Pions and 0 for Kaons.
At p= 5 GeV/c, Nph per unit length= 0.8 /cm for Pions and 0 for Kaons.
D b / b = tan2 q / (2 * sqrt (Nph) )

## Slide 21
Differential Cherenkov Detectors
Discovery of anti-proton
in 1955 by Chamberlain,
Segre et. al. at Berkeley.
Nobel Prize in 1959
With Solid (quartz) radiator
Very small acceptance in b and direction of the charged particle. (Narrow range in velocity and direction intervals ).
Mostly used for identifying particles in the beam lines.
D b/ b = (m12 – m22) /2 p 2 = tan q Dq
m1,m2 (particle masses )<< p ( momentum)
Db/ b from 0.011 to 4* 10 -6 achieved.

## Slide 22
Threshold Counters
BELLE: Threshold Cherenkov Detector
Five aerogel tiles
inside an aluminum box
lined with a white
reflector(Goretex reflector)
Performance from test- beam
p below and p above Threshold
Approx .
20 photoelectrons
per Pion detected
at 3.5 GeV/c
More than 3s
separation

## Slide 23
RICH Detectors
Measures both the Cherenkov angle and the number of photoelectrons detected.
Can be used over particle identification over large surfaces.
Requires photodetectors with single photon identification capability.

## Slide 24
RICH detectors
RICH detectors have better resolution than equivalent Differential and Threshold counters.
Let u= sin2 (q) = 1- (1/n2) - (m/p*n) 2
Number of standard deviations to discriminate between mass m1 an m2
= N s = (u2-u1) / ( s u * sqrt( N)) where s u : D q converted into the parameter u.
( D q = error in single photon q measurement)
At momentum p (=b E), p= sqrt((m22-m12)/(2* K * Ns )) , for b ~1
This equation can be used in the design of the RICH detectors.
One the first large size RICH detector: in DELPHI at LEP.
D b / b = tan(q) * Dqc = K where Dq c = < Dq > / sqrt(N ph ) + C
where <Dq> is the mean resolution per single photon in a ring and C is the
error contribution from the tracking , alignment etc.
For example , for 1.4 m long CF4 gas radiator at STP and a detector with N0 = 75 cm-1
K = 1.6 * 10 -6 . ( E=6.5 eV, DE = 1 eV)
This is better than similar Threshold counters by a factor 125.
This is also better than similar Differential counters by a factor 2.
Reason: RICH measures both q and Nph directly.

## Slide 25
In this lecture, we focus on some of the aspects related to the detection of photoelectrons in Cherenkov Detectors.
Gas based detectors:
MWPC (Multi Wire Proportional Chambers)
GEM (Gas Electron Multiplier )
Vacuum based detectors: PMT (Photomultiplier tubes)
HPD (Hybrid Photodiodes)
Detection of Photoelectrons
Convert Photons  Photoelectrons using a photocathode
Detect these photoelectrons using ‘charged track detectors’.
Measure the position and (/or ) time of photoelectrons in the tracking detector.
Principle:
General introduction to tracking detectors and silicon detectors is not covered in this lecture.
Solid state detectors: Silicon photomultipliers

## Slide 26
Photodetectors
Photon Conversion:
Photoelectric Effect : Photon energy to be above the ‘work function’
(Einstein : Nobel Prize in 1921).
Commercial alkaline Photocathodes: Bialkali , Trialkali (S20) , CsI etc.
Alkali metals have relatively low ‘work function’.
There are also gases where the photon conversion takes place.
Different photocathodes are efficient at different wavelength ranges.
Quantum Efficiency (QE) : Fraction of photons converted to electrons

## Slide 27
Photodetector with CsI photocathode
Used in ALICE experiment at CERN
Thickness of :
radiator = 10mm
quartz window= 5mm
MWPC gaps= 2 mm
Wire cathode pitch=2 mm
Anode pitch= 4 mm
anode diameter= 20 micron
pad size = 8*8 mm2
Proximity focussing
cause feedback: leading to loss of original signal info.
Backscatterd electron
Electron transport region (diffusion)
photon
photon
Typical gain: 10 4
Total detector area: 12 m2
Open geometry: using MWPC

## Slide 28
Recent Developments: Gas Based Photodetectors
GEM with semi-transparent Photocathode
(K-Cs-Sb)
Photon and ion feed back reduced.
Gated operation to reduce noise.
(no readout outside a ‘time window of signal’)
For now only closed geometry ( in sealed tubes):
Reduced fraction of useful area for photon
detection (Active Area Fraction) compared to open
geometry.
GEM: Gas Electron Multiplier

## Slide 29
Vacuum Based Photodetectors
HPD
MAPMT
PMTs
PMTs Commercially produced: more info in
www.sales.hamamatsu.com
Silicon detector of HPD

## Slide 30
Band gap in Silicon = 3.16eV; Typical Max Gain = 20 keV / 3.16 eV = 5000 (approx)
Signal pulse height spectrum of a 61-pixel HPD
Illuminated with Cherenkov photons
Features of HPD

## Slide 31
Features of the PMTs and HPDs
PMT:
Typical Gain of MAPMT 300 K.
Excellent time resolution: 125 ps for example
(Ex: used in underwater Cherenkov detectors).
Active area fraction: 40 % : Fraction of effective detection area.
This can be Improved with a lens, but then one may loose some
photons at the lens surface.
Recent developments: Flat panel pmts with 89 % active area fraction.
New photocathodes with >45% QE at 400 nm
HPD:
Typical gain 5K, but quite uniform across different channels.
Excellent Single photon identification capability.
Active area fraction: 35 76 %

## Slide 32
Comparison of photodetectors
Related to photon and ion feed back and high gains at high rate.
Detection in visible wavelength range (for better resolution)
Gaseous:
Issues:
Advantages:
Can operate in high magnetic field
Lower cost for large size detectors compared to vacuum based
Vacuum based:
Issues:
Sensitivity to magnetic field
Active Area Fraction
Advantages:
Can easily operate at high rate (eg. LHC rates and higher).
Operates also in visible wavelengths.
Ease of operation at remote locations: underwater, in space etc.
HPD: uniform gain over large number of tubes and small noise.
Choice of photodetector depends on the design of the Cherenkov detectors
and constraints on cost etc.
Other Types and new developments:
APD, Silicon photomultiplier, HAPD , MCP etc.

## Slide 33
Recent Developments:Silicon Photomultipliers
Photon Detection Efficiency (PDE)
for SiPM is better that of ordinary PMT.
Time resolution= ~ 100 ps.
Works in magnetic field
gain = ~ 10 6
Reducing noise levels for single photon
detection is still an issue and is being worked upon.
Array of APDs

## Slide 34
New Developments: Micro Channel Plate (MCP) Photon Detectors

## Slide 35
New Developments: Micro Channel Plates (MCP )
Measure Space and time of the hits.
Manufactured by industry ( Photonics for example).
Resolutions: Space: ~100 microns, Time: ~ 50 - 100 psec.
Short flight path of photoelectrons: Resistant to magnetic fields up to 0.8 Tesla.
Can work at 40 MHz readout rate.
Can detect single photons ( No noise from ‘first dynode’ as in MAPMT).
Fast ‘ageing’ at large luminosity (eg: LHC) is an issue, but there are some solutions.
Typical Size:
2 mm thickness,
51 mm X 51 mm active area.
10 micron pores separated by 15 microns
Chevron: 8 degree tilt : To increase
th gain and reduce ion-feed back
Gain: ~ 5 * 10 5
Typically ~ 1000 channels per MCP.

## Slide 36
Precision measurement of B-Decays and search for signals beyond standard model.
Two RICH detectors covering the particle momentum range 1100 GeV/c using
aerogel, C4F10 and CF4 gas radiators.
LHCb Experiment

## Slide 37
LHCb-RICH Specifications
(Refractive Index-1) vs. Photon Energy
RICH1: Aerogel 210 GeV/c
C4F10 < 70 GeV/c
RICH2: CF4 <100 GeV/c.
n=1.03
n=1.0014
n=1.0005
Aerogel C4F10 CF4
L 5 86 196 cm
qcmax 242 53 32 mrad
p Th 0.6 2.6 4.4 GeV/c
KTh 2.0 9.3 15.6 GeV/c
Aerogel: Rayleigh Scattering
Not used from 2015 onwards
T = A e (-C t / l4)
Aerogel Transmission
Typically: A= 0.94, C =0.0059 mm4 /cm

## Slide 38
LHCb- RICH1 SCHEMATIC
RICH1 OPTICS
Magnetic Shield
Beam Pipe
Photodetectors
Spherical Mirror
Flat Mirror
Gas Enclosure
Readout Electronics
Spherical Mirror tilted
to keep photodetectors outside acceptance
(tilt ~ 0.3 rad)

## Slide 39
RICH1 Photos
RICH1-HPDs
RICH1 mirrors

## Slide 40
Performance of LHCb RICH
From isolated Cherenkov rings from RICH1 in Real Data
Compare with the expectations plotted in page 9.

## Slide 41
LHCb-RICH resolution
C4F10
CF4
From simulation in 2011:
sDq = 1.53 mrad
From simulation in 2011:
sDq = 0.68 mrad
Single photon resolutions
Real Data 2011:
Resolution components: Chromatic : ref. index variation
Emission Point: tilt of the mirror
Pixel size: granularity of the pixel
PSF: spread of photo electron direction inside HPD

## Slide 42
LHCb: Hits on the RICH from Simulation
Red: From particles from Primary and Secondary Vertex
Blue: From secondaries and background processes (sometimes with
no reconstructed track)

## Slide 43
Pattern Recognition in Accelerator based Cherenkov Detector
Events with large number of charged tracks giving rise to several
overlapping Cherenkov Rings on the Photo detector plane.
Problem: To identify which tracks correspond to which hits and then identify
the type (e, p, p etc.) of the particle which created the tracks.
Hough Transform:
(used by ALICE
at CERN)
Project the particle direction on to the detector plane
Accumulate the distance of each hit from these projection points
in case of circular rings.
Collect the peaks in the accumulated set and associate the
corresponding hits to the tracks.
Likelihood Method:
For each of the track in the event, for a given mass hypothesis,
create photons and project them to the detector plane using the
knowledge of the geometry of the detector and its optical properties.
Repeat this for all the other tracks.
From this calculate the probability that a signal would be seen in each
pixel of the detector from all tracks.
Compare this with the observed set of photoelectron signal on the pixels,
by creating a likelihood.
Repeat all the above after changing the set of mass hypothesis of the tracks.
Find the set of mass hypothesis, which maximize the likelihood.
(used by LHCb
at CERN)
(Ref: R.Forty: Nucl. Inst. Mech. A 433 (1999) 257-261)

## Slide 44
LHCb-RICH PID performance
Real data
Simulation
Red: Kaon identification efficiency
Black: Pion mis-identification probability
PID performance in D* events (calibration channel)
General agreement between real data and simulation

## Slide 45
RICH Detectors in LHCb
Before RICH PID
After RICH PID
Without RICH PID, the is completely dominated by
Real data from 2011
RICH PID useful for physics analysis

## Slide 46: How to make a better RICH detector
Increase the number of photons
  Better photon detectors (photo-cathodes)
  Better area coverage by photon detectors
  Better optical coupling of components
Improve Cherenkov angle resolution
  Smaller pixels
  Lower chromatic error
    Detectors sensitive to green rather than blue
  Improve aberrations
    Light weight mirrors in the acceptance

## Slide 47: DIRC
Direct Internally Reflected Cherenkov light
Detectors outside the “acceptance”
Can fit in a narrow space
The standoff region is designed to maximize the transfer efficiency between the radiator and the detector.
If this region has the same index of refraction as the radiator, n1n2, the transfer efficiency is maximized and the image will emerge without reflection or refraction at the end surface.

## Slide 48: Babar DIRC
 8 ns Dt window for background rejection

## Slide 49: TORCH (i)
A time of flight detector based on Cherenkov radiation

## Slide 50: TORCH (ii)
The focusing unit converts angles to positions allowing the measurement of propagation time
Multiple photons from the same track can improve the time accuracy

## Slide 51
Energy loss by muons dE/dx : Bethe-Block Formula
Max energy in single
Collision.
Density correction
Ionisation Constant for material
K=4 p N re2 me c2

## Slide 52
dE/dx detectors: Silicon based
Bethe-Block formula ( for 2 g m/M << 1 )
FWHM= 30-35 %
For 0.2< b <0.9, dE/dx ~ (M/p)2
M,p= mass, momentum of the traversing particle
Ref: NIMA 469(2001) 311-315

## Slide 53
dE/dx Detectors : Silicon based
CMS Silicon tracker:
CMS Silicon strip wafer
Each Silicon sensor gives a dE/dX measurement.
Estimate the Most Probable Value from several (10-25)
measurements (Truncated Mean: Ignore upper 40%)

## Slide 54
dE/dx Detectors: Drift Chambers
Larger Landau fluctuations compared to those from Silicon detectors.
So many measurements needed to get the average.
BABAR Drfit Chamber:
Ref: IEEE-TNS VOL:47 , NO:6, Dec2000.

## Slide 55
Time of Flight (TOF) Detectors
Measurement of time: Using scintillators
Energy loss (dE/dx) from the charged particle= 2MeV/cm
This energy is re-emitted as optical photons in UV. (Approx 1 photon/100 eV)
Too small attenuation length and low yield.
Fluorescent material added to scintillator so that the photon re-emitted at
longer wavelengths (ex: 400 nm) longer attenuation length ( ex: 1 meter) and
high yield . These photons collected by a PMT. (0.002 pe per emitted photon)
NA 49: Heavy ion expt, Scintillator thickness=2.3 cm, time resolutions= 59ps, 95 ps
Had a TPC to measure dE/dx.

## Slide 56: Summary
The field of Particle Identification Detectors is an evolving field.
They have contributed to some of the important discoveries in High Energy Physics in the last 50 years and they continue to be a crucial part of some of the recent experiments.
The RICH detectors offer excellent Particle Identification capability for the hadrons since they can be designed to have very good single photon Cherenkov Angle resolution and large Photoelectron yield. Recent advances in photodetectors enhance the capability of these detectors.
The particle ID using dE/dx, time-of-flight and Transition Radiation detectors continue to provide Particle Identification in different experiments.
Particle identification is a crucial part of some of the Astroparticle physics experiments and long base line neutrino experiments

## Slide 57: The end
Thank you for
your attention
Any (more) questions?

## Slide 58
Photonic Crystals
Parrot (scarlet macaw) in a tropical forest
Optical micrograph of blue feather barbs
www.pnas.org/cgi/doi/10.1073/pnas.1204383109
CERN Courier Aug23, 2005
Phys. Rev. E 72, 010902 (2005)

## Slide 59
Photonic Crystals
Periodic arrangement of objects with two different refractive indices.
Lattice constants comparable to the
wavelength of light in the material
electrons in pure semi-conductor : similar features to photons in photonic crystal
Cherenkov radiation:
Physical origin is a mixture of conventional Cherenkov radiation
and Transition radiation
Photon propagation in the crystal in the form of Bloch waves
All the properties can be derived from solving Maxwell’s arXiV:0808:3519
equations for periodic lattice
There is no Cherenkov Threshold
Cherenkov cone can be forward or backward
Starting from two media with refractive indices n1 and n2,
create a new effective refractive index n3.
r=cylinder radius
a= lattice constant

## Slide 60
Possibility of particle identification with periodic media
PRL 113, 167402 (2014)
Sensitivity= q eletron – q proton
Transformation optics:
Longitudinal stretching=F=1.005 : Shift curves to right
Transverse stretching = G=10 : Increase q

## Slide 61
Cherenkov Detectors in Astro Particle Physics
Goal: Contribute to the understanding of our Universe.
Understanding production mechanism
(‘cosmic accelerators’) of HE cosmic rays ;
Study very energetic galactic / extragalactic objects :
SN remnants, microquasars, GRB, AGN,…;
Search for Dark matter (wimps)
. . .
SNR
AGN
GRB
Binary systems
Micro-quasars

## Slide 62
Astro Particle Physics
Neutral : Hence Weak interaction only
Neutrinos point back to the astrophysical production source
  Unlike photons which interact with CMB and matter…
  or protons: which also undergo deflection by magnetic fields
Neutrinos: Advantages:
Disadvantages:
Rate of arrival very low. Hence need very large detectors.
Using the Ocean , ice in Antartica etc.
Search for :
Neutrinos  muons
High energy Gamma and other Cosmic rays  Air showers
Ultra high energy Gamma ( > 10 19 eV )  Air showers

## Slide 63
Angle between the m and n direction=
Typically 1g / PMT
40 m from m axis
Importance of Timing Resolution
c in water ~ 20 cm/ns
Chromatic dispersion ~ 2 ns (40 m typ. Path)
(PMT TTS s ~1.3ns) so detector not dominant source of error
o
Measure position
and time of the hits.

## Slide 64
ANTARES Experiment in the sea.
Optical Module
Hamamatsu PMT :
Size :10 inch
Glass pressure Sphere.

## Slide 65
Design Specifications
IceCube Experiment in Antartica
Fully digital detector concept.
Number of strings – 75
Number of surface tanks – 160
Number of DOMs – 4820
Instrumented volume – 1 km3
Angular resolution of in-ice array < 1.0°
Fast timing: resolution < 5 ns DOM-to-DOM
Pulse resolution < 10 ns
Optical sens. 330 nm to 500 nm
Dynamic range - 1000 pe / 10 ns - 10,000 pe / 1 us.
Low noise: < 500 Hz background
High gain: O(107) PMT
Charge resolution: P/V > 2
Low power: 3.75 W
Ability to self-calibrate
Field-programmable HV generated internal to unit.
10000 psi external

## Slide 66
nm from CC
interactions
n e from CC or n x from NC interactions
nt  t  m
Ice Cube/AMANDA Event signatures
All signals from Cherenkov
Radiation.
Eµ=10 TeV
~300m for
>PeV t

## Slide 67
Knee
>1019 eV
1 km-2 year-1 sr-1
High Energy Cosmic Ray Spectrum.
Measure the Energy
Spectrum
Determine the Arrival
Direction distribution
etc.
Composed of
Baryons, photons,
neutrinos etc.

## Slide 68
Fluorescence →
Array of water → Cherenkov detectors
Principle of Auger Project

## Slide 69
1600 surface stations
1.5 km spacing
over 3000 square kilometers
Fluorescence Detectors:
4 Telescope enclosures,
each with 6 telescopes.

## Slide 70
Communications
antenna
Electronics
enclosure
Solar panels
Battery box
3 – nine inch
photomultiplier
tubes
Plastic tank with 12 tons of water
GPS antenna
AUGER Project: Water Cherenkov Detector
Time difference between test signals from nearby detectors ( Carmen-Miranda)
Installation of the Cherenkov detectors are continuing and data taking started.
First set of results are published

## Slide 71
Long Base Line Experients: T2K
Neutrino production: Protons incident on graphite target to create pions,
which decay into m and nm. The muons and and any
remaining protons and pions are stopped by a second layer
of graphite; but the nm pass through towards SuperK.
Beam energy: Centered around 600 MeV (to maximize the oscillation
probability into ne ).
Far-detector: underground to reduce cosmic rays.

## Slide 72
PMT-schematic
40 m high, 17 m radius.
Super-K water cherenkov detector

## Slide 73
Super-K Cherenkov signals
(in water)
b ~ 1
The leptons create Cherenkov light in the water.
(muon)

## Slide 74
Muons create sharp rings,
Electrons scatter and shower and hence create ‘fuzzy’ rings.
Super-K signals
Measure the charge and time of
each pulse. From this estimate :
(a) energy of the charged particle.
( Charge a number photons
a energy loss )
(b) position (vertex ) of the interaction.
(c ) direction of the charged particle.

## Slide 75
If n>2 some photons are always totally internally reflected for b1 tracks.
Radiator and light guide: Long, rectangular Synthetic Fused Silica (“Quartz”) bars (Spectrosil: average <n(l)>  1.473, radiation hard, homogenous, low chromatic dispersion)
Photons exit via wedge into expansion region
(filled with 6m3 pure, de-ionized water).
Pinhole imaging on PMT array (bar dimension small compared to standoff distance). (10,752 traditional PMTs ETL 9125, immersed in water, surrounded by hexagonal “light-catcher”, transit time spread ~1.5nsec, ~30mm diameter)
DIRC is a 3-D device, measuring: x, y and time of Cherenkov photons, defining qc, fc, tpropagation of photon.
DIRC PRINCIPLE

## Slide 76
Number of Cherenkov photons
per track (di-muons) vs. polar angle:
Resolution of Cherenkov angle fit
per track (di-muons):
s(Dqc) = 2.4 mrad
Track Cherenkov angle resolution is within ~10% of design
Between 20 and 60 signal photons per track.
DIRC PERFORMANCE

## Slide 77
D–  D0 p –
K– p +
Kaon selection efficiency for L K > L p
p mis-id as K
(6% comb. background)
(track in DIRC fiducial, comb. background corrected)
Kaon selection efficiency typically above 95%
with mis-ID of 2-10%
between 0.8-3GeV/c.
DIRC PERFORMANCE

## Slide 78
New Development: Focussing DIRC
Red photons arrive before blue photons
Time of Propagation= PathLength/ v group
Correct for Chromatic error from the
mesurement of time of propagation.
Ref: NIMA 595(2008)104-107

## Slide 79
Photon Detector of the CLEO-III Cherenkov detector
Gas Based Photon Detectors
QE: 33% at 150 nm
Spectral Range: 135-165 nm
TEA: Triethylamine
photon passes through the CaF2 and converts to photoelectron by ionizing a TEA molecule.
The photoelectron drifts towards and avalanches near the anode wires,
theirby inducing a charge signal on the cathode pads.

## Slide 80
Balloon Experiment:
RICH detector
CAPRICE Experiment
TMAE:
(tetrakis(dimethylamino)
ethylene)

## Slide 81
Components of a Cherenkov Detector
Main Components:
Radiator : To produce photons
Mirror/lens etc. : To help with the transport of photons
Photodetector : To detect the photons
Radiator: Any medium with a Refractive Index.
Aerogel: network of
SiO2 nano-crystals
The atmosphere, ocean are the radiators in some Astro Particle Cherenkov Detectors
g = 1/sqrt(1-b2)

## Slide 82
LHCb-RICH Design
RICH1: Aerogel L=5cm p:210 GeV/c
n=1.03 (nominal at 540 nm)
C4F10 L=85 cm p: < 70 GeV/c
n=1.0014 (nominal at 400 nm)
Upstream of LHCb Magnet
Acceptance: 25250 mrad (vertical)
300 mrad (horizontal)
Gas vessel: 2 X 3 X 1 m3
RICH2: CF4 L=196 cm p: < 100 GeV/c
n =1.0005 (nominal at 400 nm)
Downstream of LHCb Magnet
Acceptance: 15100 mrad (vertical)
120 mrad (horizontal)
Gas vessel : 100 m3

## Slide 83
LHCb RICH Data in Physics Analysis
Status in 2011
using RICH data
RICH data has been used in Physics analysis
f  K + K -

## Slide 84
Example of LHCb-RICH PERFORMANCE
Yield: Mean Number of hits per isolated
saturated track ( b ~1).
Aerogel | C4F10 | CF4
4.6 | 34.0 | 24
Single Photon Cherenkov Angle Resolutions in mrad.
Components and Overall (mrad) | Aerogel | C4F10 | CF4
Chromatic | 2.65 | 0.84 | 0.48
Emission Point | 0.34 | 0.61 | 0.27
Pixel Size | 0.59 | 0.6 | 0.19
PSF | 0.78 | 0.79 | 0.29
Overall RICH | 2.84 | 1.45 | 0.65
Overall RICH+Tracks | 2.86 | 1.50 | 0.76
Chromatic: From the variation in
refractive index.
Emission Point: Essentially from the
tilt of the mirrors.
Pixel Size: From the granularity of the
Silicon detector pixels in HPD
PSF ( Point Spread Function):
From the spread of the Photoelectron direction
as it travells inside the HPD,
(from the cross focussing in the electron optics)
Performance as seen in Simulated Data in 2010

## Slide 85
Differential Cherenkov Detectors
Very small acceptance in b and direction of the charged particle.
(Narrow range in velocity and direction intervals ).
From the Cherenkov angle (q ) determine b.
Mostly used for identifying particles in the beam lines.
Resolution that can be achieved = D b/ b = (m12 – m22) /2 p 2 = tan q Dq
m1,m2 (particle masses )<< p ( momentum)
At high momentum, to get better resolution, use gas radiators which have smaller
refractive index than solid radiators. Have long enough
radiators to get sufficient signal photons in the detector.
To compensate for Chromatic dispersion (n (Eph) ) , lens used in the path of the photons.
(DISC: Differential Isochronous self- collimating Cherenkov Counter).
Db/ b from 0.011 to 4* 10 -6 achieved.

## Slide 86
Differential Cherenkov Detectors
With a Gas radiator

## Slide 87
Threshold Cherenkov Detectors
Can be used over a large area, for Example : For secondary particles in a fixed target or
Collider experiment.
BELLE Experiment: To observe CP ciolation in B-meson decays at an electron-positron
collider.
Db/b = 2.3 * 10-5 using gas radiator.
E691 at Fermilab: To study decays of charm particles in the 1980’s
BELLE: Continues to take
Data.

## Slide 88
LHCb-RICH2 SCHEMATIC
Beam Axis-
Spherical Mirror
Flat Mirror
Photon funnel+Shielding
Central Tube
Support Structure
Mirror Support Panel
RICH2 Optics Top View
X
Z
Z
Y
X
Plane Mirrors to reduce the length of RICH2
Spherical mirror tilted to keep photodetectors
outside acceptance.(tilt=0.39 rad)

## Slide 89
LHCb- RICH2 STRUCTURE
Entrance Window
(PMI foam between two
carbon fibre epoxy Skins)
RICH2