| EPR - EMX |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Written by: | Dr B G Breeze | 05/04/20 |  |  |  |
| Amended by: |  |  |  |  |  |

IMPORTANT

Acid cleaning must only be conducted by authorized personal, who have had a full induction into working in the Spectroscopy labs and completed local training. A list of trained personal can be found in the webgroup px-epr-users.

Before beginning any experiment do a quick visual inspection of the spectrometer to check that all wave guides are attached and that there are no warning lights on the console. There should also be no sample in the spectrometer. Check that the desired cavity is in place.  If you have any issues, please contact a member of RTP staff.

Sample preparation

- The sample should be prepared for EPR in the prep lab (G71) not near the spectrometer. Gloves and Lab coats and goggles must be worn when in  the prep lab.

- How much sample you use will depend on your experiment.

  - Liquids: fill the tube so there is sufficient sample to fill the cavity (around 10 cm) (for a 2mm tube it is best to just fill the tube)

  - Powders: Generally, only a small volume of sample need be placed in the tube, 1-2mm is ideal. If you have low spin density up to 10 mm of sample can be placed in the tube

- Once the sample is ready it should be transferred into a ERP tube. (willmad  712-SQ-100M or 705-SQ-250M)

  - For liquid samples use a pipet with a fine needle available on the shelves in G71. When filling tube ensure that there are no air bubbles in the sample as this will causes issues in data acquisition and tuning. Air bubbles can sometimes be removed by gently tapping the tube.

  - For solid samples transfer using a clean spatula and funnel if require. Make sure that all the sample is at the bottom of the tube. Any powder stuck to the wall of the tube should be dislodged by gently tapping the tube.

- Replace cap on Tube.

- Clean any sample off the outside of the tube using lab tissues.

- Flame seal if required (see separate SSoW)

- Dispose of any pipettes and needles in the appropriate waste containers. Clean spatulas and funnels and return to shelves.

Sample Loading

- To load a  sample first loosen the collet nut on the top of the resonator.

- Check the collet is correct for you sample tube the tube. If correct move on to 10.

  - To change the colt fulling unscrew the collet nut and remove

  - Remove the collet.

  - Select the collet with correct aperture from the box on the spectrometer. The correct collet should slide onto your sample tube easily but be touching the tube.

  - Place the selected collet in the mouth of the cavity.

  - Replace the collet screw loosely.

- Before loading the sample check the position of the pedestal is not too high. If it is high loosen its nut and lower and retighten the nut.

- Load the sample by carefully inserting the tube into the collet. The tube should remain Vertical throughout loading.

- Slowly lower the tube until its base in located in the centre of the pedestal. This will ensure the tube is center and vertical in the cavity.

- Ensure that the sample is at the correct height. The centre of the sample should be at the centre of the cavity. (the center of the HS cavity is approximately 63mm below the top of the collet nut)

- Once in the correct position tighten the collet nut and confirm the pedestal is touching the bottom of the sample tube.

- The spectrometer is now safe to moved from standby in tune mode. It is important that at least 30 minutes are left between first taking the spectrometer out of standby and running the first measurement.

Spectrometer tuning

- Critcall couple the spectrometer, using separate SSoW.

- Once the tuning is complete it should be possible to set the microwaves to any power without the AFC or the Diode gauges moving.

.

Taking a spectrum

- Open a new experiment with the new experiment buttun

- Set the desired experiment parameters by clicking the experimental parameters button to open up the menu. The parameters will vary from sample to sample, a general an example  “Starting set” are outlined in the table for the reference but should not be assumed to be correct for every case . If unsure you’re the EMX parameter check spread sheet to sanity check your parameters.

| Parameter | Liquid (nitroxide radical) | Powder (nitroxide radical) | Powder (Cu2+ radical) |
| --- | --- | --- | --- |
| Center field (G) (adjust) | 3480 | 3480 | 3480 |
| Sweep Width (G) | 200 | 200 | 2000 |
| Attenuator dB | 16 | 30 | 30 |
| Gain | 1e4 | 1e4 | 1e4 |
| Modulation frequency (kHz) | 100 | 100 | 100 |
| Modulation Amplitude (G) | 0.5 | 1 | 5 |
| Modulation Phase () | 0 | 0 | 0 |
| Time Constant (ms) | 5.12 | 5.12 | 5.12 |
| Conversion time (ms) Resolution | 10.24 2048 | 10.24 4096 | 10.24 8192 |
| Number of scans | 4 | 4 | 14 |
| Other |  |  |  |

- Start a scan using the run button.

- Ensure that the spectrum is centres in the measurement and adjust centre field if not.

- Check tuning is still correct

- Start EPR frequency logger if available. If not make a note of the microwave frequency

- Begin the scan using the start scan button

- Ensure all parameter are fill out on the excel parameter sheet/Lab book.

- When scan is finished save the spectrum using the save icon. Always do this before adjusting anything else to avoid data being lost.

Sample Removal

- Put spectrometer back in tune mode. (mw button select tune form operation mode)

- Loosen Collet nut

- Remove EPR tube by gently pulling tube vertically up. Deviation away from vertical may snap the tube and irreparably damage the cavity.

- Put sample somewhere safe. And retighten collect

End of day.

- When you have run your last sample please put the spectrometer back in stand by.

- Remove all your samples and save all data.

NOTES:

If you have a liquid sample with very low spin concentration you may wish to consider using a extra high quality EPR tube. WS-!!-!!!!-!! they cost approximately but should have a much reduces background.