Risk Assessment Form

| Title of Risk Assessment | Bruker E580 EPR spectrometer Including Cryogenic Work |  |  | Date of assessment |  | 08/09/21 |
| --- | --- | --- | --- | --- | --- | --- |
| Department | Research Technology Platforms / Physics |  |  | Date review due |  | 08/09/23 |
| Description of Task/Process | Use of the Bruker E580 Pulsed EPR spectrometer in Millburn house room G78, Including use of cryogenic cooling. |  |  |  |  |  |
| Assessment carried out by | Ben Breeze |  |  |  |  |  |

| ##### Additional information | Users are responsible for having performed a risk assessment for their samples which must be logged at: https://warwick.ac.uk/research/rtp/spectroscopy/Safety/Sample When booking all users will provide a reference to the risk assessment for the sample used. Bookings are restricted to trained users who are recorded in the px-epr-users web group. |
| --- | --- |

| Hazards and how they may cause harm | Who may be at Risk? | Existing Control Measures | Current Risk Level (VL,L,M,H,VH) | Where current risk is M, H or VH, what additional Control Measures are required? | Action required by whom & by when? | Final Risk Level |
| --- | --- | --- | --- | --- | --- | --- |
| Mains Electricity Electric Shock and/or burns | Operator Other lab users | Equipment should not be tampered with. All equipment requires an in-date PAT test | (Major, Unlikely) Low | None | N/A | Low |
| Three-Phase Electricity Electric Shock and/or burns | Operator Other lab users | Three-phase power supply/cables should not be tampered with. Visual inspection should be regularly carried out. | (Major, Unlikely) Low | None | N/A | Low |
| Water leak Electric Shock and/or Burns | Operator Other lab users | Water cooling pipes/fitting should be inspected for leaks. Flooding sensors are located at the spectrometers, which shut off the water valve in event of a leak. Lab flooding procedure in event of a flood the affected area of the lab is to be cordoned off whilst the area is made safe and then dried out. Magnets should be flushed every 3 years to removes and cooling blockages. | (Minor, Unlikely) Very Low | None | N/A | Low |
| Water leak Electric Shock and/or Burns | Operator Other lab users |  | (Major, Unlikely) Low | None | N/A | Low |
| Microwave source X-Band (1 kW maximum at 9.8 Ghz) Burns to Skin and/or eyes | Operator Other lab users | Microwaves are contained with the system under normal operation. Wave guides not to be tampered with. All alterations to wave guide to be conducted by trained staff/operators. Microwave power should not be on if waveguide is open. Visual inspection of Waveguides prior to operation | (Serious, Unlikely) Low | None | N/A | Low |
| Microwave source Q Band (0.15 W maximum at 34 Ghz) Burns to Skin and/or Eyes | Operator Other lab users | Microwaves are contained with the system under normal operation. Wave guides not to be tampered with. All alterations to wave guide to be conducted by trained staff/operators. Microwave power should not be on if waveguide is open. Visual inspection of Waveguides prior to operation | (Serious, Unlikely) Low | None | N/A | Low |
| Magnetic field Bruising, from objects caught in field | Operator Other lab users | Magnetic items should not be held within the 5 G line All Lab users are made aware of the dangers and locations of High magnetic field during induction to the lab Signs are in place at all entrances to the lab alerting to the presence of high magnetic fields. All users, and especially users with pace-makers/other electronic health devices, should not cross the 5 G line which is the yoke of the spectrometer. | (minor, Likely) Low | None | N/A | Low |
| Magnetic field ill-health | Operator Other lab users |  | (Serious, Unlikely) Low | None | N/A | Low |
| Magnetic field failure of electronic health devices | Operator Other lab users |  | (Major, Unlikely) Low | None | N/A | Low |
| Cryogenics Inhalation or ingestion causing loss of consciousness, | Operator Other lab users | Read all relevant safety data sheets, University and Departmental Safety regulations and guidance, and pass any required safety tests. Room is fitted with Oxygen depletion warning sensors. | (Serious, Unlikely) Low | None | N/A | Low |
| Cryogenics Inhalation or ingestion causing asphyxiation | Operator Other lab users |  | (Extreme, Unlikely) Moderate | None | N/A | Moderate |
| Cryogenics Contact resulting in cryogenic Burns | Operator Other lab users | Read all relevant safety data sheets, University and Departmental Safety regulations and guidance, and pass any required safety tests. Appropriate clothing must be worn at all time as per university guidance PPE must be worn at all time, including Goggles and Gloves which are available in the Lab. | (Serious, Unlikely) Low | None | N/A | Low |
| Explosion/implosion of cryostat/vacuum system Cuts from Shards of Glass Metal possibly to eyes | Operator Other lab users | Use of vacuum systems prohibited unless user is trained in the safe operation of the vacuum pumps. Visual inspection of equipment before operation Safety valves are to be used and checked regularly. | (Major, Unlikely) Low | None | N/A | Low |
| Broken glassware Cuts | Operator Other lab users | Take care, dispose of sharps and broken glass in the approved receptacles provided. | (minor, Likely) Low | None | N/A | Low |
| Moving parts | Operator Other lab users | Loose clothing should be avoided and long hair should be tied back. | (Minor, Possible) Low | None | N/A | Low |
| Obstructed walkways | Operator Other lab users | Do not block walkways and be aware of risk Keep work areas clean and tidy | (Minor, Possible) Low | None | N/A | Low |
| Dirt or polluted work area | Operator Other lab users | Keep work areas clean and tidy | (Minor, Possible) Low | None | N/A | Low |

Work should not be carried out until the assessment is completed and all required control measures are in place.

| Overall Final Risk Rating (Highest level in final column above) | Moderate |
| --- | --- |

| Additional Comments from Risk Assessor (e.g. funding or practical implications) |  | None |  |  |
| --- | --- | --- | --- | --- |
| Approved By |  |  | Position |  |
| Date |  |  |  |  |

Please print a copy, sign it and keep for your records

Document History

| Version | Date | Reviewer | Comments |
| --- | --- | --- | --- |
| 1 | 05/01/18 | Ben Breeze | Re reviewed when Equipment was brought under RTP Control |
| 2 | 25/01/19 | Ben Breeze | Reviewed after change of location. No changes |
| 3 | 08/09/21 | Ben Breeze | Scheduled Review Updated Format minor changes to cryogenics procedure. |

|  | Severity |  |  |  |  |  |  | Risk Level |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Likelihood | Superficial | Minor | Serious | Major | Extreme |  | Very low | Acceptable risk - no action required |
| Unlikely | Very low | Very low | Low | Low | Moderate |  | Low | Tolerable risk - further control measures not required, but status must be monitored |
| Possible | Very low | Low | Low | Moderate | High |  | Moderate | Further control measures required to reduce risk as far as is reasonably practical |
| Likely | Low | Low | Moderate | High | Very high |  | High | Urgent action required to allow activity to continue |
| Very likely | Low | Moderate | High | Very high | Very high |  | Very high | Risk intolerable - activity must cease until the risk has been reduced |
| Extremely likely | Moderate | High | Very high | Very high | Very high |  |  |  |

See ‘Matrix for risk evaluation’ for further guidance.