Risk Assessment Form

| Title of Risk Assessment | Fumigation of rooms and MSCs (Generic) |  |  | Date of assessment | <insert date> |
| --- | --- | --- | --- | --- | --- |
| Department | All Science departments |  |  | Date review due | <review date 3 years later. To review earlier following an incident or significant change> |
| Description of Task/Process | The treatment of microbiological safety cabinets (MSC) and containment level 3 (CL3) laboratories with gaseous sterilisation agents (fumigants) to inactivate viable biological agents (e.g. after a spillage incident or prior to access for maintenance). |  |  |  |  |
| Assessment carried out by | <Insert name of individual or competent persons> |  |  |  |  |

| ##### Additional information | Note that this is a generic risk assessment so should be tailored to take into consideration any additional hazards present connected with the activity, the materials used, the people involved or the environment. Guidance on fumigation is available from the HSE website (http://www.hse.gov.uk/pUbns/priced/hsg251.pdf). NB Operators are defined in this risk assessment as those persons undertaking the fumigation operation itself, and are distinct from lab users, who have conducted operations in the lab up to the point where it required fumigation to occur. Any exposure to a toxic fumigant vapour such as formaldehyde arising from an unplanned release is likely to be reportable under the Reporting of Injuries, Diseases and Dangerous Occurrences Regulations 2013 (RIDDOR). |
| --- | --- |

| Hazards and how they may cause harm | Who may be at Risk? | Existing Control Measures | Current Risk Level (VL,L,M,H,VH) | Where current risk is M, H or VH, what additional Control Measures are required? | Action required by whom & by when? | Final Risk Level |
| --- | --- | --- | --- | --- | --- | --- |
| Discharge of fumigant from CL3 lab into adjacent areas | Operators, bystanders | Room sealability tests performed at six-monthly intervals. Effective periodic testing of air-handling systems to ensure neutral pressure of rooms. Operators trained in applying secondary sealing measures. Robust SOP in place. Withdrawal of access privileges except for operators. | L |  |  | Low |
| Discharge of fumigant from MSC into room | Operators, bystanders | Purchase of good quality MSCs with certified sealability. Operators trained in secondary sealing of MSC night doors, etc. Robust SOP in place to ensure effective scavenging/discharge. Laboratory closed during fumigation cycles. | L |  |  | Low |
| Leakage of fumigant from ductwork into plant rooms, etc. | Estates staff, contractors | Gas-tight dampers in place in supply and extract ductwork. Effective programme of maintenance of gas-tight dampers and other mechanical controls. Regular periodic checks of other ductwork for joint leaks, etc. Permits to access plant rooms not issued during fumigation cycles. | L |  |  | Low |
| Exposure to fumigant during release via chimney stack | Estates staff, contractors | Chimney stacks of sufficient height to ensure free venting of fumigant. Permits to access relevant roof areas not issued during fumigation cycles. | L |  |  | Low |
| Inhalation of fumigant vapour while conducting fumigation | Operators | Training of operators in correct procedure. Use of calibrated formaldemeter to identify unsafe conditions. Robust SOP in place. Use of fit-tested respiratory protective equipment. | L |  |  | Low |
| Damage to skin/eyes from liquid fumigant | Operators | Training of operators in correct procedure. Robust SOP in place. Use of personal protective equipment (eye protection, gloves and labcoat). | L |  |  | Low |
| Infection by exposure to viable biological agents | Lab users, operators, Bystanders, contractors | Establishment of negative air pressure regime. Use of terminal HEPA filtration on exhaust air ducts. Use of primary containment equipment (i.e. MSCs). Robust SOP in place for fumigation cycle. Fumigation cycle validated to inactivate relevant organisms. Fumigation following a spillage incident is operable without access to CL3 laboratory. | L |  |  | Low |

Work should not be carried out until the assessment is completed and all required control measures are in place.

| Overall Final Risk Rating (Highest level in final column above) | L |
| --- | --- |

| Additional Comments from Risk Assessor (e.g. funding or practical implications) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Approved By | <Name of peer, supervisor or line manager> |  | Position | <insert role> |
| Date | <insert date of approval> |  |  |  |

Please print a copy, sign it and keep for your records

Document history:

| Version | Date | Reviewer | Comments/Changes Made |
| --- | --- | --- | --- |
| V1.0 | 12/12/2017 | Ian Graham | New version |
| V1.1 | 15/08/2019 | Ian Graham | Reviewed – no significant changes made |
| V1.2 | 03/08/2023 | Caroline Farren | Reviewed – added document history and altered dates on front page – this section should be completed by those adopting this generic. |

|  | Severity |  |  |  |  |  |  | Risk Level |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Likelihood | Superficial | Minor | Serious | Major | Extreme |  | Very low | Acceptable risk - no action required |
| Unlikely | Very low | Very low | Low | Low | Moderate |  | Low | Tolerable risk - further control measures not required, but status must be monitored |
| Possible | Very low | Low | Low | Moderate | High |  | Moderate | Further control measures required to reduce risk as far as is reasonably practical |
| Likely | Low | Low | Moderate | High | Very high |  | High | Urgent action required to allow activity to continue |
| Very likely | Low | Moderate | High | Very high | Very high |  | Very high | Risk intolerable - activity must cease until the risk has been reduced |
| Extremely likely | Moderate | High | Very high | Very high | Very high |  |  |  |

See ‘Matrix for risk evaluation’ for further guidance.