Risk Assessment Form

| Title of Risk Assessment | Handling cryogenic fluids |  |  | Date of assessment |  | <insert date> |
| --- | --- | --- | --- | --- | --- | --- |
| Department | All Science Departments |  |  | Date review due |  | <review date 3 years later. To review earlier following an incident or significant change> |
| Description of Task/Process | The filling, dispensing and use of cryogenic fluids |  |  |  |  |  |
| Assessment carried out by | <name of individual or competent persons> |  |  |  |  |  |

| ##### Additional information | Note that this is a generic risk assessment so should be tailored to take into consideration any additional hazards present connected with the activity, the materials used, the people involved or the environment. This risk assessment should be read in conjunction with Substance Information Note 018 which contains general information from the safety data sheet (SDS). |
| --- | --- |

| Hazards and how they may cause harm | Who may be at Risk? | Existing Control Measures | Current Risk Level (VL,L,M,H,VH) | Where current risk is M, H or VH, what additional Control Measures are required? | Action required by whom & by when? | Final Risk Level |
| --- | --- | --- | --- | --- | --- | --- |
| Cold burns to hands from fluid (filling and use) | User | - PPE - cryogenic gloves must be worn. BS EN 420, BS EN 511, BS EN 388 - Training to highlight this need. - Signage. | M | - Training to be carried out by competent person. - Signage to be put into place where absent in cryogenic dispensing areas. | - Line manager to ensure that users are trained prior to permitting them to handle cryogens. - Technical Services to be requested to procure and fix up relevant hazard signage. | Low |
| Cryogenic splash in eyes (filling and use) | User, bystander | - Cryogenics handled in restricted area or an area which is located away from other activities. - PPE – eye protection to be worn BS EN 166 grade 3, goggles or full-face visor - Training to highlight this need. - Signage. | L | Suitability of cryogenic handling to be considered in future building design. | Estates Project Managers | Low |
| Cryogenic burns to feet (filling and use) | User, bystander | - Enclosed footwear must be worn. If boots are worn, trousers should be worn outside the boots to prevent liquid running into the top of the boot. - Training to highlight this need. | L |  |  | Low |
| Cryogenic burns caused by failure of protection devices | User, bystander | - Regular servicing and replacement of protective devices (pressure relief valves, bursting discs) | L | - Pressurised dewars and liquid gas tanks to be subject to regular servicing and statutory testing if not already in place. Note records need to be maintained to evidence this. - New facilities to be added to Crimson system by Department | Technical Services to ensure regular PPM in place and that records are maintained and the installation and pressure dewars are added to Crimson. | Low |
| Asphyxiation when filling | User, bystander | - Fill in well ventilated area, preferably outside. - Facilities suitably designed with adequate ventilation. - Remain attentive to filling operation and do not overfill. - Indoor filling points to have oxygen level sensor system and alarm. - Alarm system subject to PPM and testing. | H | - All persons involved in the filling, dispensing and use of cryogenic fluids to know what to do in the event of an alarm sounding. - Persons NOT to ignore alarm. - If alarm continually sounds in area, report the matter to Technical Services and the H&S Department. - Servicing regime to be in place to ensure alarm system remains operational at all times. | - Technical Services and Supervisors to ensure that persons are trained prior to handling cryogenics. - Technical Services to ensure PPM in place for system alarm and to maintain records. | M |
| Cryogenic burns from contact with cold surfaces | User | PPE - cryogenic gloves to BS EN 420, BS EN 511, BS EN 388. Training to alert user to this hazard. | L |  |  | Low |
| Injury to back, arms, legs when moving trolley (slips, trips) | User, bystander | - Dewars have wheels that can aid transport from one location to another. - Manual handling training. Awareness of push rather than pull. Awareness of uneven surfaces likely to cause topple. - Route to be pre-planned, to avoid slopes and stairs. | M | Designers of new facilities to consider the movement of dewars etc as part of the design scope to ensure slopes and steps are eliminated. | Estates Project Managers and Department involved in design of facility. | Low |
| Asphyxiation when using or transporting | User, bystander | - Work in well-ventilated area. If working in an enclosed space use oxygen depletion alarm. - Suitable procedures in place where there is a need to transport filled dewar in a lift to prevent others from using lift at the same time. | M | Persons to be trained in emergency procedures if alarm sounds. | Local manager / supervisor to ensure persons are trained in the emergency procedures prior to any work taking place involving cryogens. | Low |
| Obstruction of walkways – fire escape issues | User, bystander | Users are aware that they need to maintain thoroughfares at all times. | L | Regular checks of thoroughfares to be carried out by departmental staff. To report issues as they arise. | Any person in the Department (at all times) | Low |

Work should not be carried out until the assessment is completed and all required control measures are in place.

| Overall Final Risk Rating (Highest level in final column above) | M |
| --- | --- |

| Additional Comments from Risk Assessor (e.g. funding or practical implications) |  |  |  |  |
| --- | --- | --- | --- | --- |
| Approved By | <name of peer, supervisor or line manager> |  | Position | <insert role> |
| Date | <insert date of approval> |  |  |  |

Please print a copy, sign it and keep for your records

Document history:

| Version | Date | Reviewer | Comments/Changes Made |
| --- | --- | --- | --- |
| V1.0 | 15/07/2017 | Caroline Farren | New version |
| V2.0 | 15/08/2019 | Caroline Farren | Reviewed – no significant changes made |
| V2.1 | 03/08/2023 | Caroline Farren | Reviewed – added document history and altered dates on front page – this section should be completed by those adopting this generic. |

|  | Severity |  |  |  |  |  |  | Risk Level |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Likelihood | Superficial | Minor | Serious | Major | Extreme |  | Very low | Acceptable risk - no action required |
| Unlikely | Very low | Very low | Low | Low | Moderate |  | Low | Tolerable risk - further control measures not required, but status must be monitored |
| Possible | Very low | Low | Low | Moderate | High |  | Moderate | Further control measures required to reduce risk as far as is reasonably practical |
| Likely | Low | Low | Moderate | High | Very high |  | High | Urgent action required to allow activity to continue |
| Very likely | Low | Moderate | High | Very high | Very high |  | Very high | Risk intolerable - activity must cease until the risk has been reduced |
| Extremely likely | Moderate | High | Very high | Very high | Very high |  |  |  |

See ‘Matrix for risk evaluation’ for further guidance.