## Slide 1
Medical device to generate pressure
The design, research, development and testing of the device
Introduction
In the medical field, the need for a reusable, low cost, handheld device to generate negative pressure using a syringe has been identified. This project was based upon the design of such a device. The design thought process can be seen below:
Syringe
Easy insertion
For a variety of sizes and types
Multiple Use
Sterilisable in an autoclave i.e. resistant to sterilising liquid and high heat
Hand-Held
Ergonomics
Smooth action
Low Cost
Low cost material
One system
Single-person & single hand
Able to create suitable pressure
Longevity
Low Cost
Multiple Use
Use of Syringe
CONSTRAINTS
Conclusion
The manufactured prototype was not fully able to fulfill the performance as predicted by the MATLAB® model. Performance was hindered by insufficient spring strength in the automated plunger action and by the cracking pressure of the valves used. Nonetheless a maximum pressure of
-73kPa was produced when bypassing the automated plunger action.
From these results, only minor modifications were necessary to be made so as to reach our aim. However, after much analysis of the prototype, the design was deemed to make syringe insertion too slow, meaning more augmentation. Several improved prototype designs have been created, although a single, optimum and all encompassing design has yet to be been found.
The project still has scope for further development, yet the contribution of the work undertaken thus far, has been vital to advancing this project ever closer to completion.
Objective
To design, test and build a working device built to create a target pressure between -60kPa and -80kPa. The design had to take in account the constraints to make the optimum design. The initial prototypes included ball valves and springs to create an automated system. However, due to patent pending issues, the design and specific use cannot be publicly shown.
Daniel Lester:
Mechanical Engineering 3rd Year
d.j.lester@warwick.ac.uk
Dr Duncan Lockerby (Engineering Supervisor)
Duncan.lockerby@warwick.ac.uk
1
Use MATLAB® to create a theoretical model, describing the relationship between the physical parameters of the device and the incremental changes in pressure in the system, to verify the devices capabilities.
2
Testing of ideas for design, including:
3
Finalise Design
Research into CE Marking to allow the medical device to be freely marketed in the EU without further control.
Building test models for verification of concepts
Efficiency of springs (used with valve)
Contact with Prototype manufacturers
Effectiveness of Rapid Prototype model and valves
4. Testing
After much communication with various local manufacturers a prototype was fabricated by CLA Tools, fitted with check valves from Lee Products Ltd and springs from Irvine Springs. Carrying out tests allowed comparisons to the MATLAB® model (left).
The prototype results show a considerable difference from the MATLAB® model due to assumptions required to simplify the mathematics. These include the frictional within the syringe, as well as within the prototype design.
The graphs shows that the springs for the automated plunger action greatly reduce the capability of the device. This is likely to be due to their spring constant value being too low.
Both MATLAB® models use the same code but with a different value for diameter of chamber 2 (V2). This was found to be the deciding factor in achieving maximum pressure. The MATLAB® model of the new design shows a reduction in V2 diameter and subsequently an increase in maximum achievable negative pressure.
6. CE Marking (EU Medical Device Regulations)
The device is a Class I medical device. Subsequently, in order for it to bear the CE mark, it must follow procedure of Annex VII of Medical Devices Regulations directive 93/42. Annex VII states that ‘the manufacturer’ must keep records of: necessary technical drawing and calculations, results of appropriate risk analysis, test reports and clinical data from clinical evaluations.
Clinical evaluations are used to verify proper device performance and determine undesirable side-effects and must be performed ethically and in a manner which guarantees scientific validity of conclusions.
In order to perform clinical trials in the EU the device must follow the procedure of Annex VIII of directive 93/42. Annex VIII states that ‘the manufacturer' must provide: an investigation plan, the opinion of the ethics committee concerned, the name of the medical practitioner responsible for the investigations and a statement that the device in question conforms to the essential requirements and that every precaution has been taken to protect the health and safety of the patient.
Paul Davis:
General Engineering 3rd Year
p.l.davis@warwick.ac.uk
3. Theory and Mathematical Modelling
By applying basic scientific principles from the mathematics of pressure systems to our specific design it was possible to derive the equations for the model. MATLAB® calculated results for the position of the syringe plunger and the pressure achieved after each pump cycle of the syringe, based on these equations.
For n=2:11, Solve roots to find Pn (for first cycle P1 = Patm), Use Pn to find xn, Repeat for each pump cycle.
Position of plunger after cycle ‘n’
Volume of chamber 3
Volume of chamber 2
Cross-sectional area of syringe
Spring constant
Atmospheric Pressure
Pressure after cycle ‘n’
Maximum plunger compression
Chart: Pump Cycle Number
Prototype (no springs) | 29.78 | 45.4833 | 53.1167 | 55.56 | 63.26 | 66.29 | 67.6467 | 68.6267 | 69.32 | 69.5133
Prototype (with springs) | 18.1367 | 28.2567 | 34.1233 | 37.39 | 39.7167 | 40.9533 | 41.39 | 41.6433 | 41.7333 | 41.57
Test Model | 25.45 | 44.0467 | 53.8367 | 60.6167 | 64.85 | 67.2567 | 68.7333 | 69.67 | 70.0933 | 70.48
Matlab Model of Prototype | 66.2 | 73.82 | 74.64 | 74.72 | 74.73 | 74.73 | 74.73 | 74.73 | 74.73 | 74.73
Matlab Model of new design | 72.2 | 81.58 | 82.73 | 82.88 | 82.89 | 82.9 | 82.9 | 82.9 | 82.9 | 82.9
The required pressure is generated by the pumping action of the syringe drawing air from V3 into V2 through a one-way valve and expelling it through another one-way valve to the atmosphere. This decreases air density in V3 and thus increases negative pressure
Comparison Between Theoretical and Experimental Pressure Values